## Summary & Method
Robotic Surgical Instrument Adverse Events — FDA MAUDE Analysis
Product code NAY (System, Surgical, Computer Controlled Instrument: da Vinci system, EndoWrist instruments & accessories). Window 2021–2026 (date_received).
Source: openFDA device/event (MAUDE), pulled 2026-06-25, 107,279 reports. Denominators: Intuitive preliminary FY results (worldwide da Vinci procedures).
KEY FINDINGS
1. Raw report totals are distorted. | Total reports rose ~25x (2.0k→49.3k); normalised, all-reports/100k procedures still rose ~12.5x — a reporting/recall artifact, NOT a real failure surge.
2. The artifact is 'Material Split, Cut or Torn'. | Concentrated in high-volume reusable instruments (fenestrated bipolar, monopolar curved scissors, ProGrasp); these have very low harm fractions (~2%).
3. Patient-harm rate is flat-to-declining. | Injury+Death per 100k procedures: 36 (2021) → ~74–77 (2022–23) → 61 (2025). Death/100k stable at ~2.6–5.3.
4. Harm concentrates in active energy instruments. | Harmonic, Tip Cover Accessory, Vessel Sealer, SynchroSeal, Force Bipolar carry high harm fractions (10–38%).
5. Component DETACHMENT is the dominant harm mode. | #1 failure mode for nearly every instrument; Harmonic alone has 728 detachment-coded harm reports.
6. Device-linked deaths sit with hemostasis devices. | Vessel Sealer (15 deaths) and Stapler (6) — seal/staple-line failure → haemorrhage. System 'deaths' (cart/arm 352) are mostly 'no device problem identified' = procedure deaths.
CAVEATS
Passive surveillance | Under-reporting, duplicate/supplemental MDRs, unverified causality. Counts are REPORTS, not validated events.
Event type | 'Injury'/'Death' is the harm signal; 'Malfunction' alone is not. Severity is sometimes reclassified on manufacturer follow-up.
Denominator | Worldwide procedures used; MAUDE is US-FDA but includes manufacturer-reported OUS events. A US-only rate (see toggle) lifts levels ~1.3–1.5x, same shape.
Instrument normalisation | No per-instrument procedure denominator exists; instrument tables are absolute counts/shares, not per-procedure rates.
2026 | Partial year (through ~June); excluded from annual rate calcs.

## Normalised Rates
Procedure-normalised reporting & harm rates
Blue = hardcoded input. Black = formula. Yellow = assumption you can change.
US share of da Vinci procedures (estimate): |  |  |  | 70% | <- adjust; US-only rates (cols K–M) scale by this
Year | Procedures (WW) | All reports | Injury | Death | Injury+Death | All /100k | Harm /100k | Death /100k | US proc (est) | US all/100k | US harm/100k | Source (proc)
2021 | 1594000 | 1997 | 532 | 43 | 575 | 125.2823 | 36.0728 | 2.6976 | 1115800 | 178.9747 | 51.5325 | FY2021 prelim
2022 | 1875000 | 5727 | 1275 | 100 | 1375 | 305.44 | 73.3333 | 5.3333 | 1312500 | 436.3429 | 104.7619 | FY2022 prelim
2023 | 2286000 | 10898 | 1694 | 63 | 1757 | 476.7279 | 76.8591 | 2.7559 | 1600200 | 681.0399 | 109.7988 | FY2023 prelim
2024 | 2683000 | 12872 | 1777 | 69 | 1846 | 479.7615 | 68.8036 | 2.5717 | 1878100 | 685.3735 | 98.2908 | FY2024 prelim
2025 | 3153000 | 49306 | 1833 | 98 | 1931 | 1563.7805 | 61.2433 | 3.1082 | 2207100 | 2233.9722 | 87.4904 | FY2025 prelim
2026 | 1313750

## Instrument League Table
Failure & harm by instrument (full set, 2021–2026)
Reports, Injury+Death and harm fraction split by report year (2021–26; 2026 partial). Reports per surgical case = year reports ÷ that year's da Vinci procedures (linked to 'Normalised Rates'). Total reports = 2021–26 sum. Harm fraction = (Injury+Death)/Reports.
Instrument family | Group | Total reports 2021–26 | Reports by year | [5 empty] | Injury+Death by year | [5 empty] | Harm fraction by year | [5 empty] | Reports per 1000 surgical cases by year |  |  |  |  | Top failure modes (Injury+Death)
|  |  | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2021 | 2022 | 2023 | 2024 | 2025
Patient-Side Cart/Arm | system | 9741 | 181 | 1251 | 1736 | 1517 | 2103 | 2953 | 137 | 479 | 386 | 475 | 508 | 281 | 75.69% | 38.29% | 22.24% | 31.31% | 24.16% | 9.516% | 11.36% | 66.72% | 75.94% | 56.54% | 66.7% | Adverse Event Without Identified Device or Use Problem (1554); Visual Prompts will not Clear (195); Insufficient Device Problem Information (191); Unintended System Motion (47)
Fenestrated Bipolar | energy | 18497 | 201 | 415 | 993 | 1421 | 10756 | 4711 | 16 | 51 | 79 | 145 | 111 | 43 | 7.96% | 12.29% | 7.956% | 10.2% | 1.032% | 0.9128% | 12.61% | 22.13% | 43.44% | 52.96% | 341.1% | Detachment of Device or Device Component (154); Material Fragmentation (145); Break (39); Material Split, Cut or Torn (18)
Monopolar Curved Scissors | energy | 12812 | 31 | 165 | 295 | 495 | 7914 | 3912 | 18 | 39 | 50 | 76 | 70 | 30 | 58.06% | 23.64% | 16.95% | 15.35% | 0.8845% | 0.7669% | 1.945% | 8.8% | 12.9% | 18.45% | 251% | Detachment of Device or Device Component (82); Adverse Event Without Identified Device or Use Problem (39); Material Fragmentation (38); Material Separation (22)
Vision System | system | 6164 | 33 | 614 | 1043 | 1227 | 1597 | 1650 | 7 | 53 | 92 | 96 | 146 | 57 | 21.21% | 8.632% | 8.821% | 7.824% | 9.142% | 3.455% | 2.07% | 32.75% | 45.63% | 45.73% | 50.65% | Visual Prompts will not Clear (108); Output Problem (74); Adverse Event Without Identified Device or Use Problem (48); Energy Output Problem (48)
ProGrasp Forceps | mechanical | 9876 | 36 | 113 | 241 | 296 | 6189 | 3001 | 11 | 23 | 46 | 74 | 49 | 19 | 30.56% | 20.35% | 19.09% | 25% | 0.7917% | 0.6331% | 2.258% | 6.027% | 10.54% | 11.03% | 196.3% | Detachment of Device or Device Component (78); Material Fragmentation (67); Material Split, Cut or Torn (20); Entrapment of Device (14)
Maryland Bipolar | energy | 8108 | 109 | 257 | 691 | 664 | 4636 | 1751 | 8 | 26 | 30 | 29 | 34 | 14 | 7.339% | 10.12% | 4.342% | 4.367% | 0.7334% | 0.7995% | 6.838% | 13.71% | 30.23% | 24.75% | 147% | Detachment of Device or Device Component (37); Material Fragmentation (29); Adverse Event Without Identified Device or Use Problem (14); Break (9)
Generic System/Platform | system | 2605 | 315 | 280 | 351 | 823 | 564 | 272 | 50 | 48 | 47 | 95 | 75 | 29 | 15.87% | 17.14% | 13.39% | 11.54% | 13.3% | 10.66% | 19.76% | 14.93% | 15.35% | 30.67% | 17.89% | Adverse Event Without Identified Device or Use Problem (154); Insufficient Device Problem Information (28); Detachment of Device or Device Component (18); Break (17)
Needle Driver / SutureCut | mechanical | 7374 | 33 | 124 | 337 | 1403 | 3587 | 1890 | 23 | 49 | 51 | 59 | 54 | 24 | 69.7% | 39.52% | 15.13% | 4.205% | 1.505% | 1.27% | 2.07% | 6.613% | 14.74% | 52.29% | 113.8% | Detachment of Device or Device Component (148); Material Fragmentation (61); Break (20); Difficult to Open or Close (12)
Force Bipolar | energy | 4310 | 61 | 177 | 537 | 617 | 1860 | 1058 | 18 | 53 | 96 | 105 | 117 | 35 | 29.51% | 29.94% | 17.88% | 17.02% | 6.29% | 3.308% | 3.827% | 9.44% | 23.49% | 23% | 58.99% | Detachment of Device or Device Component (119); Entrapment of Device (86); Material Fragmentation (76); Difficult to Open or Close (47)
Clip Applier | mechanical | 4744 | 59 | 328 | 868 | 1008 | 1720 | 761 | 7 | 44 | 74 | 78 | 57 | 19 | 11.86% | 13.41% | 8.525% | 7.738% | 3.314% | 2.497% | 3.701% | 17.49% | 37.97% | 37.57% | 54.55% | Detachment of Device or Device Component (130); Difficult to Open or Close (108); Material Separation (17); Unintended System Motion (10)
Cadiere Forceps | mechanical | 2822 | 37 | 62 | 183 | 81 | 1589 | 870 | 22 | 27 | 60 | 15 | 22 | 13 | 59.46% | 43.55% | 32.79% | 18.52% | 1.385% | 1.494% | 2.321% | 3.307% | 8.005% | 3.019% | 50.4% | Detachment of Device or Device Component (73); Material Fragmentation (25); Adverse Event Without Identified Device or Use Problem (15); Break (14)
Tip-Up Fenestrated Grasper | mechanical | 2303 | 11 | 79 | 131 | 77 | 1179 | 826 | 8 | 11 | 19 | 20 | 24 | 9 | 72.73% | 13.92% | 14.5% | 25.97% | 2.036% | 1.09% | 0.6901% | 4.213% | 5.731% | 2.87% | 37.39% | Detachment of Device or Device Component (31); Material Fragmentation (30); Break (9); Adverse Event Without Identified Device or Use Problem (7)
Bipolar (other) | energy | 1255 | 16 | 37 | 120 | 91 | 678 | 313 | 6 | 7 | 18 | 10 | 16 | 4 | 37.5% | 18.92% | 15% | 10.99% | 2.36% | 1.278% | 1.004% | 1.973% | 5.249% | 3.392% | 21.5% | Detachment of Device or Device Component (22); Material Fragmentation (13); Material Split, Cut or Torn (8); Adverse Event Without Identified Device or Use Problem (7)
Harmonic (ultrasonic) | energy | 2687 | 68 | 258 | 758 | 487 | 661 | 455 | 66 | 147 | 311 | 210 | 229 | 68 | 97.06% | 56.98% | 41.03% | 43.12% | 34.64% | 14.95% | 4.266% | 13.76% | 33.16% | 18.15% | 20.96% | Detachment of Device or Device Component (728); Material Fragmentation (163); Break (116); Material Separation (19)
Surgeon Console | system | 2548 | 93 | 383 | 570 | 456 | 579 | 467 | 2 | 25 | 45 | 47 | 38 | 20 | 2.151% | 6.527% | 7.895% | 10.31% | 6.563% | 4.283% | 5.834% | 20.43% | 24.93% | 17% | 18.36% | Visual Prompts will not Clear (38); No Display/Image (28); Adverse Event Without Identified Device or Use Problem (21); No Device Output (11)
Vessel Sealer | energy | 2077 | 104 | 234 | 417 | 370 | 652 | 300 | 40 | 85 | 91 | 103 | 113 | 46 | 38.46% | 36.32% | 21.82% | 27.84% | 17.33% | 15.33% | 6.524% | 12.48% | 18.24% | 13.79% | 20.68% | Failure to Deliver Energy (134); Failure to Seal (79); Material Fragmentation (62); Detachment of Device or Device Component (50)
Permanent Cautery Hook | energy | 1777 | 108 | 150 | 282 | 304 | 645 | 288 | 14 | 19 | 44 | 39 | 64 | 25 | 12.96% | 12.67% | 15.6% | 12.83% | 9.922% | 8.681% | 6.775% | 8% | 12.34% | 11.33% | 20.46% | Detachment of Device or Device Component (96); Material Fragmentation (63); Break (15); Arcing (12)
Tenaculum Forceps | mechanical | 1421 | 114 | 133 | 212 | 364 | 464 | 134 | 1 | 10 | 7 | 9 | 10 | 0 | 0.8772% | 7.519% | 3.302% | 2.473% | 2.155% | 0% | 7.152% | 7.093% | 9.274% | 13.57% | 14.72% | Material Fragmentation (16); Detachment of Device or Device Component (7); Break (5); Difficult to Remove (4)
Grasping Retractor | mechanical | 1106 | 116 | 79 | 136 | 301 | 376 | 98 | 6 | 3 | 4 | 13 | 5 | 1 | 5.172% | 3.797% | 2.941% | 4.319% | 1.33% | 1.02% | 7.277% | 4.213% | 5.949% | 11.22% | 11.93% | Detachment of Device or Device Component (19); Material Fragmentation (8); Material Split, Cut or Torn (2); Break (1)
Endoscope/Camera | system | 955 | 11 | 91 | 220 | 166 | 314 | 153 | 0 | 8 | 9 | 8 | 4 | 3 | 0% | 8.791% | 4.091% | 4.819% | 1.274% | 1.961% | 0.6901% | 4.853% | 9.624% | 6.187% | 9.959% | Display or Visual Feedback Problem (6); Output Problem (3); Misfocusing (3); Material Fragmentation (2)
Permanent Cautery Spatula | energy | 774 | 66 | 79 | 108 | 111 | 285 | 125 | 5 | 11 | 10 | 15 | 12 | 6 | 7.576% | 13.92% | 9.259% | 13.51% | 4.211% | 4.8% | 4.141% | 4.213% | 4.724% | 4.137% | 9.039% | Detachment of Device or Device Component (24); Material Fragmentation (16); Adverse Event Without Identified Device or Use Problem (5); Break (4)
Tip Cover Accessory | energy | 821 | 59 | 93 | 151 | 181 | 216 | 121 | 39 | 45 | 63 | 35 | 90 | 39 | 66.1% | 48.39% | 41.72% | 19.34% | 41.67% | 32.23% | 3.701% | 4.96% | 6.605% | 6.746% | 6.851% | Material Separation (147); Detachment of Device or Device Component (108); Arcing (24); Material Split, Cut or Torn (18)
Stapler (SureForm/EndoWrist) | mechanical | 937 | 60 | 206 | 230 | 139 | 197 | 105 | 24 | 56 | 50 | 32 | 28 | 13 | 40% | 27.18% | 21.74% | 23.02% | 14.21% | 12.38% | 3.764% | 10.99% | 10.06% | 5.181% | 6.248% | Difficult to Open or Close (28); Detachment of Device or Device Component (27); Adverse Event Without Identified Device or Use Problem (24); Failure to Form Staple (21)
Curved Bipolar Dissector | energy | 303 | 5 | 8 | 31 | 12 | 173 | 74 | 1 | 1 | 2 | 0 | 2 | 2 | 20% | 12.5% | 6.452% | 0% | 1.156% | 2.703% | 0.3137% | 0.4267% | 1.356% | 0.4473% | 5.487% | Detachment of Device or Device Component (3); Material Fragmentation (3); Adverse Event Without Identified Device or Use Problem (1); Break (1)
SynchroSeal | energy | 614 | 49 | 57 | 131 | 123 | 167 | 87 | 32 | 32 | 36 | 27 | 28 | 17 | 65.31% | 56.14% | 27.48% | 21.95% | 16.77% | 19.54% | 3.074% | 3.04% | 5.731% | 4.584% | 5.297% | Detachment of Device or Device Component (82); Material Fragmentation (25); Failure to Deliver Energy (12); Material Separation (10)
Monopolar Cautery (other) | energy | 193 | 2 | 16 | 26 | 46 | 77 | 26 | 0 | 0 | 1 | 1 | 1 | 1 | 0% | 0% | 3.846% | 2.174% | 1.299% | 3.846% | 0.1255% | 0.8533% | 1.137% | 1.714% | 2.442% | Material Fragmentation (2); Use of Device Problem (1); Thermal Decomposition of Device (1); Detachment of Device or Device Component (1)
Suction Irrigator | accessory | 176 | 1 | 13 | 48 | 43 | 48 | 23 | 1 | 9 | 20 | 15 | 12 | 6 | 100% | 69.23% | 41.67% | 34.88% | 25% | 26.09% | 0.06274% | 0.6933% | 2.1% | 1.603% | 1.522% | Detachment of Device or Device Component (43); Material Fragmentation (10); Difficult to Remove (4); Suction Problem (3)
Grasper/Forceps (other) | mechanical | 104 | 4 | 5 | 14 | 15 | 45 | 21 | 2 | 3 | 3 | 0 | 3 | 0 | 50% | 60% | 21.43% | 0% | 6.667% | 0% | 0.2509% | 0.2667% | 0.6124% | 0.5591% | 1.427% | Detachment of Device or Device Component (8); Mechanical Problem (1); Break (1); Material Fragmentation (1)
Cannula/Trocar/Reducer/Sheath | accessory | 100 | 11 | 13 | 12 | 31 | 22 | 11 | 11 | 9 | 11 | 14 | 6 | 8 | 100% | 69.23% | 91.67% | 45.16% | 27.27% | 72.73% | 0.6901% | 0.6933% | 0.5249% | 1.155% | 0.6977% | Detachment of Device or Device Component (29); Material Fragmentation (9); Insufficient Device Problem Information (7); Adverse Event Without Identified Device or Use Problem (4)
Cobra Grasper | mechanical | 35 | 2 | 5 | 6 | 3 | 11 | 8 | 0 | 1 | 0 | 1 | 2 | 0 | 0% | 20% | 0% | 33.33% | 18.18% | 0% | 0.1255% | 0.2667% | 0.2625% | 0.1118% | 0.3489% | Entrapment of Device (2); Defective Device (1); Material Fragmentation (1)
Unknown/Unspecified | unknown | 17 | 0 | 0 | 0 | 0 | 2 | 15 | 0 | 0 | 0 | 0 | 1 | 0 |  |  |  |  | 50% | 0% | 0% | 0% | 0% | 0% | 0.06343% | Adverse Event Without Identified Device or Use Problem (1)
PK Dissecting Forceps | energy | 23 | 1 | 2 | 20 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0% | 50% | 10% |  |  |  | 0.06274% | 0.1067% | 0.8749% | 0% | 0% | Break (2); Insufficient Device Problem Information (1); Detachment of Device or Device Component (1)

## Instrument x Year (harm)
Injury+Death by instrument and year
2025 now complete; 2026 partial (omitted).
family | 2021 | 2022 | 2023 | 2024 | 2025
Fenestrated Bipolar | 16 | 51 | 79 | 145 | 111
Monopolar Curved Scissors | 18 | 39 | 50 | 76 | 70
ProGrasp Forceps | 11 | 23 | 46 | 74 | 49
Patient-Side Cart/Arm | 137 | 479 | 386 | 475 | 508
Maryland Bipolar | 8 | 26 | 30 | 29 | 34
Needle Driver / SutureCut | 23 | 49 | 51 | 59 | 54
Vision System | 7 | 53 | 92 | 96 | 146
Clip Applier | 7 | 44 | 74 | 78 | 57
Force Bipolar | 18 | 53 | 96 | 105 | 117
Cadiere Forceps | 22 | 27 | 60 | 15 | 22
Harmonic (ultrasonic) | 66 | 147 | 311 | 210 | 229
Generic System/Platform | 50 | 48 | 47 | 95 | 75
Surgeon Console | 2 | 25 | 45 | 47 | 38
Tip-Up Fenestrated Grasper | 8 | 11 | 19 | 20 | 24
Vessel Sealer | 40 | 85 | 91 | 103 | 113
Permanent Cautery Hook | 14 | 19 | 44 | 39 | 64
Tenaculum Forceps | 1 | 10 | 7 | 9 | 10
Bipolar (other) | 6 | 7 | 18 | 10 | 16
Grasping Retractor | 6 | 3 | 4 | 13 | 5
Endoscope/Camera | 0 | 8 | 9 | 8 | 4
Stapler (SureForm/EndoWrist) | 24 | 56 | 50 | 32 | 28
Tip Cover Accessory | 39 | 45 | 63 | 35 | 90
Permanent Cautery Spatula | 5 | 11 | 10 | 15 | 12
SynchroSeal | 32 | 32 | 36 | 27 | 28
Curved Bipolar Dissector | 1 | 1 | 2 | 0 | 2
Monopolar Cautery (other) | 0 | 0 | 1 | 1 | 1
Suction Irrigator | 1 | 9 | 20 | 15 | 12
Grasper/Forceps (other) | 2 | 3 | 3 | 0 | 3
Cannula/Trocar/Reducer/Sheath | 11 | 9 | 11 | 14 | 6
Cobra Grasper | 0 | 1 | 0 | 1 | 2
PK Dissecting Forceps | 0 | 1 | 2 | 0 | 0
Unknown/Unspecified | 0 | 0 | 0 | 0 | 1

## Failure x Instrument
Device-problem code x instrument (Injury+Death)
family | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem | Material Fragmentation | Insufficient Device Problem Information | Visual Prompts will not Clear | Break | Material Separation | Difficult to Open or Close | Failure to Deliver Energy | Entrapment of Device | Material Split, Cut or Torn | Output Problem | Energy Output Problem | Difficult to Remove | Unintended System Motion | No Device Output | Arcing | Failure to Seal | Appropriate Device Problem Term/ Code Not Available | Thermal Decomposition of Device | Defective Device | No Display/ Image | Unintended Movement | Physical Resistance/Sticking | Use of Device Problem
Fenestrated Bipolar | 154 | 13 | 145 | 8 | 0 | 39 | 9 | 11 | 8 | 11 | 18 | 1 | 13 | 10 | 3 | 2 | 9 | 0 | 3 | 1 | 2 | 0 | 0 | 0 | 1
Monopolar Curved Scissors | 82 | 39 | 38 | 10 | 0 | 17 | 22 | 3 | 0 | 3 | 11 | 0 | 3 | 7 | 2 | 0 | 20 | 0 | 2 | 6 | 4 | 0 | 4 | 1 | 5
ProGrasp Forceps | 78 | 2 | 67 | 4 | 0 | 7 | 6 | 7 | 0 | 14 | 20 | 1 | 0 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
Patient-Side Cart/Arm | 12 | 1554 | 6 | 191 | 195 | 4 | 2 | 8 | 1 | 3 | 2 | 36 | 4 | 5 | 47 | 47 | 6 | 2 | 29 | 0 | 3 | 5 | 15 | 22 | 8
Maryland Bipolar | 37 | 14 | 29 | 5 | 0 | 9 | 0 | 4 | 4 | 4 | 4 | 0 | 5 | 6 | 1 | 3 | 5 | 0 | 1 | 7 | 1 | 0 | 0 | 2 | 2
Needle Driver / SutureCut | 148 | 2 | 61 | 1 | 0 | 20 | 9 | 12 | 0 | 0 | 8 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 0
Vision System | 0 | 48 | 0 | 25 | 108 | 1 | 1 | 0 | 24 | 0 | 0 | 74 | 48 | 1 | 4 | 8 | 4 | 0 | 9 | 2 | 0 | 7 | 0 | 0 | 0
Clip Applier | 130 | 7 | 9 | 8 | 0 | 3 | 17 | 108 | 0 | 2 | 0 | 0 | 0 | 3 | 10 | 1 | 0 | 0 | 2 | 0 | 6 | 0 | 7 | 0 | 0
Force Bipolar | 119 | 8 | 76 | 5 | 0 | 28 | 9 | 47 | 2 | 86 | 21 | 0 | 2 | 36 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 0 | 1 | 1 | 0
Cadiere Forceps | 73 | 15 | 25 | 2 | 0 | 14 | 4 | 6 | 0 | 4 | 8 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 3 | 0 | 2 | 0 | 1 | 0 | 0
Harmonic (ultrasonic) | 728 | 3 | 163 | 0 | 4 | 116 | 19 | 1 | 1 | 0 | 0 | 5 | 2 | 0 | 3 | 9 | 0 | 4 | 0 | 15 | 6 | 0 | 0 | 1 | 1
Generic System/Platform | 18 | 154 | 11 | 28 | 1 | 17 | 6 | 4 | 1 | 1 | 0 | 5 | 5 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 5 | 0 | 2 | 0 | 5
Surgeon Console | 0 | 21 | 0 | 3 | 38 | 0 | 0 | 4 | 1 | 0 | 0 | 9 | 2 | 0 | 8 | 11 | 0 | 0 | 2 | 0 | 0 | 28 | 5 | 3 | 4
Tip-Up Fenestrated Grasper | 31 | 7 | 30 | 4 | 0 | 9 | 1 | 1 | 0 | 3 | 1 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0
Vessel Sealer | 50 | 30 | 62 | 25 | 1 | 9 | 4 | 37 | 134 | 17 | 3 | 0 | 27 | 6 | 4 | 1 | 1 | 79 | 8 | 2 | 7 | 0 | 0 | 6 | 0
Permanent Cautery Hook | 96 | 3 | 63 | 1 | 0 | 15 | 8 | 1 | 1 | 0 | 6 | 0 | 2 | 4 | 1 | 0 | 12 | 0 | 0 | 5 | 2 | 0 | 1 | 0 | 1
Tenaculum Forceps | 7 | 0 | 16 | 0 | 0 | 5 | 0 | 1 | 0 | 1 | 3 | 0 | 0 | 4 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Bipolar (other) | 22 | 7 | 13 | 0 | 0 | 3 | 3 | 0 | 1 | 2 | 8 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Grasping Retractor | 19 | 0 | 8 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
Endoscope/Camera | 2 | 0 | 2 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0
Stapler (SureForm/EndoWrist) | 27 | 24 | 14 | 11 | 0 | 3 | 7 | 28 | 0 | 4 | 2 | 3 | 0 | 8 | 4 | 7 | 0 | 0 | 1 | 0 | 1 | 0 | 2 | 0 | 1
Tip Cover Accessory | 108 | 3 | 8 | 2 | 0 | 0 | 147 | 0 | 0 | 0 | 18 | 0 | 1 | 0 | 0 | 0 | 24 | 0 | 1 | 15 | 1 | 0 | 0 | 0 | 0
Permanent Cautery Spatula | 24 | 5 | 16 | 0 | 0 | 4 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
SynchroSeal | 82 | 10 | 25 | 6 | 0 | 5 | 10 | 2 | 12 | 2 | 3 | 0 | 1 | 0 | 2 | 0 | 6 | 5 | 0 | 1 | 1 | 0 | 0 | 0 | 0
Curved Bipolar Dissector | 3 | 1 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Monopolar Cautery (other) | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1
Suction Irrigator | 43 | 1 | 10 | 2 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Grasper/Forceps (other) | 8 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cannula/Trocar/Reducer/Sheath | 29 | 4 | 9 | 7 | 0 | 3 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3
Cobra Grasper | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
PK Dissecting Forceps | 1 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unknown/Unspecified | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Failure Mode Deep Dive
Column Labels
| 2021 |  | 2022 |  | 2023 |  | 2024 |  | 2025 |  | 2026 |  | Total Neither_ | Total Injury_
Row Labels | Neither_ | Injury_ | Neither_ | Injury_ | Neither_ | Injury_ | Neither_ | Injury_ | Neither_ | Injury_ | Neither_ | Injury_
Cadiere Forceps | 14 | 21 | 23 | 16 | 133 | 54 | 51 | 10 | 1206 | 16 | 559 | 8 | 1986 | 125
Detachment of Device or Device Component | 10 | 19 | 8 | 11 | 57 | 35 | 13 | 5 | 37 | 2 | 15 | 1 | 140 | 73
Material Fragmentation | 1 | 0 | 4 | 2 | 15 | 10 | 9 | 3 | 81 | 7 | 4 | 3 | 114 | 25
Material Separation |  |  | 0 | 1 | 4 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 5 | 4
Break | 3 | 2 | 4 | 0 | 52 | 8 | 14 | 2 | 38 | 2 | 28 | 0 | 139 | 14
Material Split, Cut or Torn |  |  | 0 | 1 |  |  | 2 | 0 | 1030 | 4 | 499 | 3 | 1531 | 8
Unintended System Motion |  |  | 7 | 1 | 5 | 0 | 12 | 0 | 20 | 0 | 13 | 0 | 57 | 1
Clip Applier | 36 | 3 | 76 | 27 | 83 | 50 | 100 | 41 | 590 | 37 | 270 | 11 | 1155 | 169
Detachment of Device or Device Component | 16 | 3 | 17 | 20 | 14 | 40 | 14 | 31 | 29 | 30 | 11 | 6 | 101 | 130
Material Fragmentation |  |  | 11 | 3 | 14 | 1 | 28 | 1 | 41 | 3 | 5 | 1 | 99 | 9
Material Separation |  |  | 1 | 1 | 2 | 6 | 1 | 3 | 0 | 3 | 0 | 4 | 4 | 17
Break |  |  | 1 | 0 | 12 | 0 | 12 | 2 | 40 | 1 | 27 | 0 | 92 | 3
Material Split, Cut or Torn |  |  | 1 | 0 | 3 | 0 | 5 | 0 | 455 | 0 | 214 | 0 | 678 | 0
Unintended System Motion | 20 | 0 | 45 | 3 | 38 | 3 | 40 | 4 | 25 | 0 | 13 | 0 | 181 | 10
Fenestrated Bipolar | 161 | 13 | 221 | 34 | 538 | 65 | 857 | 129 | 8954 | 95 | 3491 | 32 | 14222 | 368
Detachment of Device or Device Component | 5 | 10 | 33 | 26 | 96 | 41 | 61 | 54 | 143 | 19 | 88 | 4 | 426 | 154
Material Fragmentation | 4 | 0 | 27 | 2 | 127 | 14 | 282 | 50 | 956 | 61 | 38 | 18 | 1434 | 145
Material Separation | 2 | 0 | 2 | 0 | 7 | 1 | 3 | 4 | 0 | 1 | 0 | 3 | 14 | 9
Break | 9 | 2 | 7 | 3 | 93 | 7 | 214 | 20 | 269 | 4 | 152 | 3 | 744 | 39
Material Split, Cut or Torn | 141 | 1 | 139 | 2 | 176 | 2 | 261 | 1 | 7516 | 9 | 3182 | 3 | 11415 | 18
Unintended System Motion |  |  | 13 | 1 | 39 | 0 | 36 | 0 | 70 | 1 | 31 | 1 | 189 | 3
Grasper/Forceps (other) | 1 | 1 | 2 | 3 | 8 | 3 | 13 | 0 | 36 | 3 | 15 | 0 | 75 | 10
Detachment of Device or Device Component | 1 | 1 | 1 | 3 | 4 | 2 | 1 | 0 | 3 | 2 | 2 | 0 | 12 | 8
Material Fragmentation |  |  | 1 | 0 | 1 | 0 | 9 | 0 | 5 | 1 | 1 | 0 | 17 | 1
Break |  |  |  |  | 0 | 1 |  |  | 1 | 0 | 2 | 0 | 3 | 1
Material Split, Cut or Torn |  |  |  |  | 1 | 0 |  |  | 26 | 0 | 10 | 0 | 37 | 0
Unintended System Motion |  |  |  |  | 2 | 0 | 3 | 0 | 1 | 0 |  |  | 6 | 0
Grasping Retractor | 112 | 7 | 63 | 3 | 116 | 4 | 280 | 11 | 285 | 4 | 45 | 1 | 901 | 30
Detachment of Device or Device Component | 98 | 6 | 36 | 2 | 32 | 4 | 15 | 5 | 14 | 1 | 2 | 1 | 197 | 19
Material Fragmentation | 2 | 1 | 24 | 1 | 53 | 0 | 242 | 5 | 50 | 1 |  |  | 371 | 8
Material Separation | [6 empty] | 1 | 0 |  |  |  |  | 1 | 0
Break | 8 | 0 |  |  | 6 | 0 | 11 | 1 | 9 | 0 |  |  | 34 | 1
Material Split, Cut or Torn | 4 | 0 | 3 | 0 | 25 | 0 | 9 | 0 | 208 | 2 | 43 | 0 | 292 | 2
Unintended System Motion | [6 empty] | 2 | 0 | 4 | 0 |  |  | 6 | 0
Maryland Bipolar | 61 | 5 | 111 | 7 | 292 | 14 | 303 | 20 | 3554 | 26 | 1176 | 8 | 5497 | 80
Detachment of Device or Device Component | 3 | 4 | 10 | 6 | 41 | 12 | 29 | 10 | 72 | 3 | 30 | 2 | 185 | 37
Material Fragmentation | 0 | 1 | 18 | 0 | 65 | 1 | 74 | 8 | 314 | 14 | 17 | 5 | 488 | 29
Material Separation |  |  | 2 | 0 | 4 | 0 | 2 | 0 |  |  |  |  | 8 | 0
Break | 4 | 0 | 6 | 1 | 49 | 1 | 73 | 2 | 135 | 4 | 51 | 1 | 318 | 9
Material Split, Cut or Torn | 54 | 0 | 65 | 0 | 97 | 0 | 108 | 0 | 2998 | 4 | 1056 | 0 | 4378 | 4
Unintended System Motion |  |  | 10 | 0 | 36 | 0 | 17 | 0 | 35 | 1 | 22 | 0 | 120 | 1
Monopolar Curved Scissors | 7 | 11 | 65 | 8 | 176 | 30 | 274 | 48 | 6769 | 50 | 3252 | 25 | 10543 | 172
Detachment of Device or Device Component | 1 | 8 | 21 | 6 | 64 | 23 | 56 | 18 | 72 | 18 | 70 | 9 | 284 | 82
Material Fragmentation | 0 | 1 | 12 | 1 | 29 | 2 | 49 | 13 | 767 | 15 | 86 | 6 | 943 | 38
Material Separation | 0 | 1 | 1 | 0 | 3 | 2 | 0 | 2 | 2 | 11 | 1 | 6 | 7 | 22
Break |  |  | 6 | 1 | 28 | 2 | 86 | 13 | 222 | 1 | 158 | 0 | 500 | 17
Material Split, Cut or Torn | 6 | 1 | 11 | 0 | 14 | 0 | 48 | 2 | 5670 | 4 | 2918 | 4 | 8667 | 11
Unintended System Motion |  |  | 14 | 0 | 38 | 1 | 35 | 0 | 36 | 1 | 19 | 0 | 142 | 2
Needle Driver / SutureCut | 11 | 25 | 57 | 46 | 265 | 51 | 1291 | 56 | 2832 | 50 | 1342 | 19 | 5798 | 247
Detachment of Device or Device Component | 10 | 22 | 23 | 39 | 59 | 40 | 32 | 26 | 64 | 14 | 30 | 7 | 218 | 148
Material Fragmentation |  |  | 13 | 2 | 68 | 4 | 1115 | 25 | 266 | 24 | 13 | 6 | 1475 | 61
Material Separation | 0 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 4 | 1 | 0 | 5 | 9
Break | 1 | 2 | 6 | 3 | 37 | 6 | 57 | 4 | 102 | 3 | 48 | 2 | 251 | 20
Material Split, Cut or Torn |  |  | 2 | 0 | 17 | 0 | 47 | 0 | 2368 | 4 | 1228 | 4 | 3662 | 8
Unintended System Motion |  |  | 12 | 0 | 83 | 0 | 38 | 0 | 32 | 1 | 22 | 0 | 187 | 1
Permanent Cautery Hook | 67 | 16 | 81 | 16 | 210 | 38 | 185 | 35 | 417 | 61 | 167 | 23 | 1127 | 189
Detachment of Device or Device Component | 39 | 13 | 49 | 12 | 71 | 28 | 40 | 22 | 30 | 17 | 16 | 4 | 245 | 96
Material Fragmentation | 0 | 1 | 11 | 2 | 43 | 3 | 59 | 10 | 31 | 34 | 9 | 13 | 153 | 63
Material Separation | 1 | 0 | 3 | 0 | 3 | 1 | 1 | 0 | 0 | 3 | 0 | 4 | 8 | 8
Break | 12 | 1 | 4 | 2 | 65 | 4 | 48 | 3 | 103 | 4 | 47 | 1 | 279 | 15
Material Split, Cut or Torn | 15 | 1 | 13 | 0 | 13 | 1 | 21 | 0 | 234 | 3 | 91 | 1 | 387 | 6
Unintended System Motion |  |  | 1 | 0 | 15 | 1 | 16 | 0 | 19 | 0 | 4 | 0 | 55 | 1
ProGrasp Forceps | 8 | 7 | 45 | 20 | 145 | 34 | 171 | 62 | 4578 | 40 | 1969 | 15 | 6916 | 178
Detachment of Device or Device Component | 7 | 4 | 21 | 15 | 61 | 26 | 61 | 28 | 136 | 4 | 80 | 1 | 366 | 78
Material Fragmentation | 1 | 3 | 9 | 5 | 12 | 3 | 55 | 26 | 555 | 25 | 12 | 5 | 644 | 67
Material Separation |  |  |  |  | 4 | 2 | 4 | 3 |  |  | 0 | 1 | 8 | 6
Break |  |  | 7 | 0 | 34 | 2 | 31 | 4 | 96 | 1 | 52 | 0 | 220 | 7
Material Split, Cut or Torn |  |  | 2 | 0 | 3 | 1 | 3 | 1 | 3747 | 10 | 1805 | 8 | 5560 | 20
Unintended System Motion |  |  | 6 | 0 | 31 | 0 | 17 | 0 | 44 | 0 | 20 | 0 | 118 | 0
Tip Cover Accessory | 17 | 43 | 27 | 36 | 59 | 52 | 97 | 30 | 96 | 82 | 69 | 38 | 365 | 281
Detachment of Device or Device Component | 0 | 13 | 0 | 11 | 12 | 34 | 13 | 16 | 11 | 23 | 9 | 11 | 45 | 108
Material Fragmentation | 1 | 1 | 0 | 2 | 2 | 1 | 3 | 1 | 1 | 2 | 1 | 1 | 8 | 8
Material Separation | 0 | 23 | 0 | 21 | 8 | 13 | 4 | 11 | 11 | 55 | 3 | 24 | 26 | 147
Break |  |  | 1 | 0 | 5 | 0 | 10 | 0 | 9 | 0 | 9 | 0 | 34 | 0
Material Split, Cut or Torn | 16 | 6 | 26 | 2 | 32 | 4 | 67 | 2 | 64 | 2 | 47 | 2 | 252 | 18
Grand Total | 495 | 152 | 771 | 216 | 2025 | 395 | 3622 | 442 | 29317 | 464 | 12355 | 180 | 48585 | 1849

## Failure Mode Deep Dive (2)
Failure-mode deep dive — by year, instrument, mode and harm category
Tidy/long format: one row per failure-mode x instrument x year x harm category, with report count. A report listing several problem codes is counted under each (so totals across modes exceed unique reports). Harm category: Death / Injury / Neither (malfunction, other or unspecified). 2026 partial. Use the column filters to slice.
Filtered report-count total: | 63170
Failure mode | Instrument family | Group | Year | Harm category | Neither | Death | Injury | Neither% | Death% | Injury% | Reports/1000 surgeries | Report count
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2021 | Injury | 0 | 0 | 61 | 0% | 0% | 3.827% | 3.827% | 61
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2022 | Injury | 0 | 0 | 128 | 0% | 0% | 6.827% | 6.827% | 128
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2022 | Neither | 88 | 0 | 0 | 4.693% | 0% | 0% | 4.693% | 88
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2023 | Injury | 0 | 0 | 253 | 0% | 0% | 11.07% | 11.07% | 253
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2023 | Neither | 278 | 0 | 0 | 12.16% | 0% | 0% | 12.16% | 278
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2024 | Injury | 0 | 0 | 125 | 0% | 0% | 4.659% | 4.659% | 125
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2024 | Neither | 116 | 0 | 0 | 4.324% | 0% | 0% | 4.324% | 116
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2025 | Injury | 0 | 0 | 123 | 0% | 0% | 3.901% | 3.901% | 123
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2025 | Neither | 46 | 0 | 0 | 1.459% | 0% | 0% | 1.459% | 46
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2026 | Injury | 0 | 0 | 38 | 0% | 0% | 2.892% | 2.892% | 38
Detachment of Device or Device Component | Harmonic (ultrasonic) | energy | 2026 | Neither | 25 | 0 | 0 | 1.903% | 0% | 0% | 1.903% | 25
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2021 | Injury | 0 | 0 | 10 | 0% | 0% | 0.6274% | 0.6274% | 10
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2021 | Neither | 5 | 0 | 0 | 0.3137% | 0% | 0% | 0.3137% | 5
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2022 | Injury | 0 | 0 | 26 | 0% | 0% | 1.387% | 1.387% | 26
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2022 | Neither | 33 | 0 | 0 | 1.76% | 0% | 0% | 1.76% | 33
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2023 | Injury | 0 | 0 | 41 | 0% | 0% | 1.794% | 1.794% | 41
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2023 | Neither | 96 | 0 | 0 | 4.199% | 0% | 0% | 4.199% | 96
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2024 | Injury | 0 | 0 | 54 | 0% | 0% | 2.013% | 2.013% | 54
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2024 | Neither | 61 | 0 | 0 | 2.274% | 0% | 0% | 2.274% | 61
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 19 | 0% | 0% | 0.6026% | 0.6026% | 19
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2025 | Neither | 143 | 0 | 0 | 4.535% | 0% | 0% | 4.535% | 143
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Detachment of Device or Device Component | Fenestrated Bipolar | energy | 2026 | Neither | 88 | 0 | 0 | 6.698% | 0% | 0% | 6.698% | 88
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2021 | Injury | 0 | 0 | 4 | 0% | 0% | 0.2509% | 0.2509% | 4
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2021 | Neither | 7 | 0 | 0 | 0.4391% | 0% | 0% | 0.4391% | 7
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2022 | Injury | 0 | 0 | 15 | 0% | 0% | 0.8% | 0.8% | 15
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2022 | Neither | 21 | 0 | 0 | 1.12% | 0% | 0% | 1.12% | 21
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2023 | Injury | 0 | 0 | 26 | 0% | 0% | 1.137% | 1.137% | 26
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2023 | Neither | 61 | 0 | 0 | 2.668% | 0% | 0% | 2.668% | 61
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2024 | Injury | 0 | 0 | 28 | 0% | 0% | 1.044% | 1.044% | 28
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2024 | Neither | 61 | 0 | 0 | 2.274% | 0% | 0% | 2.274% | 61
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2025 | Neither | 136 | 0 | 0 | 4.313% | 0% | 0% | 4.313% | 136
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | ProGrasp Forceps | mechanical | 2026 | Neither | 80 | 0 | 0 | 6.089% | 0% | 0% | 6.089% | 80
Detachment of Device or Device Component | Force Bipolar | energy | 2021 | Injury | 0 | 0 | 6 | 0% | 0% | 0.3764% | 0.3764% | 6
Detachment of Device or Device Component | Force Bipolar | energy | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Detachment of Device or Device Component | Force Bipolar | energy | 2022 | Injury | 0 | 0 | 23 | 0% | 0% | 1.227% | 1.227% | 23
Detachment of Device or Device Component | Force Bipolar | energy | 2022 | Neither | 18 | 0 | 0 | 0.96% | 0% | 0% | 0.96% | 18
Detachment of Device or Device Component | Force Bipolar | energy | 2023 | Injury | 0 | 0 | 38 | 0% | 0% | 1.662% | 1.662% | 38
Detachment of Device or Device Component | Force Bipolar | energy | 2023 | Neither | 69 | 0 | 0 | 3.018% | 0% | 0% | 3.018% | 69
Detachment of Device or Device Component | Force Bipolar | energy | 2024 | Injury | 0 | 0 | 34 | 0% | 0% | 1.267% | 1.267% | 34
Detachment of Device or Device Component | Force Bipolar | energy | 2024 | Neither | 87 | 0 | 0 | 3.243% | 0% | 0% | 3.243% | 87
Detachment of Device or Device Component | Force Bipolar | energy | 2025 | Injury | 0 | 0 | 14 | 0% | 0% | 0.444% | 0.444% | 14
Detachment of Device or Device Component | Force Bipolar | energy | 2025 | Neither | 96 | 0 | 0 | 3.045% | 0% | 0% | 3.045% | 96
Detachment of Device or Device Component | Force Bipolar | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Detachment of Device or Device Component | Force Bipolar | energy | 2026 | Neither | 46 | 0 | 0 | 3.501% | 0% | 0% | 3.501% | 46
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2021 | Injury | 0 | 0 | 22 | 0% | 0% | 1.38% | 1.38% | 22
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2021 | Neither | 10 | 0 | 0 | 0.6274% | 0% | 0% | 0.6274% | 10
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2022 | Injury | 0 | 0 | 39 | 0% | 0% | 2.08% | 2.08% | 39
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2022 | Neither | 23 | 0 | 0 | 1.227% | 0% | 0% | 1.227% | 23
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2023 | Injury | 0 | 0 | 40 | 0% | 0% | 1.75% | 1.75% | 40
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2023 | Neither | 59 | 0 | 0 | 2.581% | 0% | 0% | 2.581% | 59
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2024 | Injury | 0 | 0 | 26 | 0% | 0% | 0.9691% | 0.9691% | 26
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2024 | Neither | 32 | 0 | 0 | 1.193% | 0% | 0% | 1.193% | 32
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 14 | 0% | 0% | 0.444% | 0.444% | 14
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2025 | Neither | 64 | 0 | 0 | 2.03% | 0% | 0% | 2.03% | 64
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2026 | Injury | 0 | 0 | 7 | 0% | 0% | 0.5328% | 0.5328% | 7
Detachment of Device or Device Component | Needle Driver / SutureCut | mechanical | 2026 | Neither | 30 | 0 | 0 | 2.284% | 0% | 0% | 2.284% | 30
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2021 | Injury | 0 | 0 | 8 | 0% | 0% | 0.5019% | 0.5019% | 8
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2022 | Injury | 0 | 0 | 6 | 0% | 0% | 0.32% | 0.32% | 6
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2022 | Neither | 21 | 0 | 0 | 1.12% | 0% | 0% | 1.12% | 21
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2023 | Injury | 0 | 0 | 23 | 0% | 0% | 1.006% | 1.006% | 23
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2023 | Neither | 64 | 0 | 0 | 2.8% | 0% | 0% | 2.8% | 64
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2024 | Injury | 0 | 0 | 18 | 0% | 0% | 0.6709% | 0.6709% | 18
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2024 | Neither | 56 | 0 | 0 | 2.087% | 0% | 0% | 2.087% | 56
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 18 | 0% | 0% | 0.5709% | 0.5709% | 18
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2025 | Neither | 72 | 0 | 0 | 2.284% | 0% | 0% | 2.284% | 72
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2026 | Injury | 0 | 0 | 9 | 0% | 0% | 0.6851% | 0.6851% | 9
Detachment of Device or Device Component | Monopolar Curved Scissors | energy | 2026 | Neither | 70 | 0 | 0 | 5.328% | 0% | 0% | 5.328% | 70
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2021 | Injury | 0 | 0 | 13 | 0% | 0% | 0.8156% | 0.8156% | 13
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2021 | Neither | 39 | 0 | 0 | 2.447% | 0% | 0% | 2.447% | 39
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2022 | Injury | 0 | 0 | 12 | 0% | 0% | 0.64% | 0.64% | 12
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2022 | Neither | 49 | 0 | 0 | 2.613% | 0% | 0% | 2.613% | 49
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 28 | 0% | 0% | 1.225% | 1.225% | 28
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2023 | Neither | 71 | 0 | 0 | 3.106% | 0% | 0% | 3.106% | 71
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2024 | Injury | 0 | 0 | 22 | 0% | 0% | 0.82% | 0.82% | 22
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2024 | Neither | 40 | 0 | 0 | 1.491% | 0% | 0% | 1.491% | 40
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2025 | Injury | 0 | 0 | 17 | 0% | 0% | 0.5392% | 0.5392% | 17
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2025 | Neither | 30 | 0 | 0 | 0.9515% | 0% | 0% | 0.9515% | 30
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Detachment of Device or Device Component | Permanent Cautery Hook | energy | 2026 | Neither | 16 | 0 | 0 | 1.218% | 0% | 0% | 1.218% | 16
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2021 | Neither | 105 | 0 | 0 | 6.587% | 0% | 0% | 6.587% | 105
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2022 | Neither | 87 | 0 | 0 | 4.64% | 0% | 0% | 4.64% | 87
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2023 | Neither | 79 | 0 | 0 | 3.456% | 0% | 0% | 3.456% | 79
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2024 | Neither | 31 | 0 | 0 | 1.155% | 0% | 0% | 1.155% | 31
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2025 | Neither | 9 | 0 | 0 | 0.2854% | 0% | 0% | 0.2854% | 9
Detachment of Device or Device Component | Tenaculum Forceps | mechanical | 2026 | Neither | 5 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 5
Detachment of Device or Device Component | Clip Applier | mechanical | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Detachment of Device or Device Component | Clip Applier | mechanical | 2021 | Neither | 16 | 0 | 0 | 1.004% | 0% | 0% | 1.004% | 16
Detachment of Device or Device Component | Clip Applier | mechanical | 2022 | Injury | 0 | 0 | 20 | 0% | 0% | 1.067% | 1.067% | 20
Detachment of Device or Device Component | Clip Applier | mechanical | 2022 | Neither | 17 | 0 | 0 | 0.9067% | 0% | 0% | 0.9067% | 17
Detachment of Device or Device Component | Clip Applier | mechanical | 2023 | Injury | 0 | 0 | 40 | 0% | 0% | 1.75% | 1.75% | 40
Detachment of Device or Device Component | Clip Applier | mechanical | 2023 | Neither | 14 | 0 | 0 | 0.6124% | 0% | 0% | 0.6124% | 14
Detachment of Device or Device Component | Clip Applier | mechanical | 2024 | Injury | 0 | 0 | 31 | 0% | 0% | 1.155% | 1.155% | 31
Detachment of Device or Device Component | Clip Applier | mechanical | 2024 | Neither | 14 | 0 | 0 | 0.5218% | 0% | 0% | 0.5218% | 14
Detachment of Device or Device Component | Clip Applier | mechanical | 2025 | Injury | 0 | 0 | 30 | 0% | 0% | 0.9515% | 0.9515% | 30
Detachment of Device or Device Component | Clip Applier | mechanical | 2025 | Neither | 29 | 0 | 0 | 0.9198% | 0% | 0% | 0.9198% | 29
Detachment of Device or Device Component | Clip Applier | mechanical | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Detachment of Device or Device Component | Clip Applier | mechanical | 2026 | Neither | 11 | 0 | 0 | 0.8373% | 0% | 0% | 0.8373% | 11
Detachment of Device or Device Component | Maryland Bipolar | energy | 2021 | Injury | 0 | 0 | 4 | 0% | 0% | 0.2509% | 0.2509% | 4
Detachment of Device or Device Component | Maryland Bipolar | energy | 2021 | Neither | 3 | 0 | 0 | 0.1882% | 0% | 0% | 0.1882% | 3
Detachment of Device or Device Component | Maryland Bipolar | energy | 2022 | Injury | 0 | 0 | 6 | 0% | 0% | 0.32% | 0.32% | 6
Detachment of Device or Device Component | Maryland Bipolar | energy | 2022 | Neither | 10 | 0 | 0 | 0.5333% | 0% | 0% | 0.5333% | 10
Detachment of Device or Device Component | Maryland Bipolar | energy | 2023 | Injury | 0 | 0 | 12 | 0% | 0% | 0.5249% | 0.5249% | 12
Detachment of Device or Device Component | Maryland Bipolar | energy | 2023 | Neither | 41 | 0 | 0 | 1.794% | 0% | 0% | 1.794% | 41
Detachment of Device or Device Component | Maryland Bipolar | energy | 2024 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3727% | 0.3727% | 10
Detachment of Device or Device Component | Maryland Bipolar | energy | 2024 | Neither | 29 | 0 | 0 | 1.081% | 0% | 0% | 1.081% | 29
Detachment of Device or Device Component | Maryland Bipolar | energy | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Detachment of Device or Device Component | Maryland Bipolar | energy | 2025 | Neither | 72 | 0 | 0 | 2.284% | 0% | 0% | 2.284% | 72
Detachment of Device or Device Component | Maryland Bipolar | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Detachment of Device or Device Component | Maryland Bipolar | energy | 2026 | Neither | 30 | 0 | 0 | 2.284% | 0% | 0% | 2.284% | 30
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2021 | Injury | 0 | 0 | 6 | 0% | 0% | 0.3764% | 0.3764% | 6
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2021 | Neither | 98 | 0 | 0 | 6.148% | 0% | 0% | 6.148% | 98
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2022 | Neither | 36 | 0 | 0 | 1.92% | 0% | 0% | 1.92% | 36
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2023 | Neither | 32 | 0 | 0 | 1.4% | 0% | 0% | 1.4% | 32
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2024 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1864% | 0.1864% | 5
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2024 | Neither | 15 | 0 | 0 | 0.5591% | 0% | 0% | 0.5591% | 15
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2025 | Neither | 14 | 0 | 0 | 0.444% | 0% | 0% | 0.444% | 14
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Grasping Retractor | mechanical | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2021 | Injury | 0 | 0 | 19 | 0% | 0% | 1.192% | 1.192% | 19
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2021 | Neither | 10 | 0 | 0 | 0.6274% | 0% | 0% | 0.6274% | 10
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2022 | Injury | 0 | 0 | 11 | 0% | 0% | 0.5867% | 0.5867% | 11
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2022 | Neither | 8 | 0 | 0 | 0.4267% | 0% | 0% | 0.4267% | 8
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2023 | Injury | 0 | 0 | 35 | 0% | 0% | 1.531% | 1.531% | 35
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2023 | Neither | 57 | 0 | 0 | 2.493% | 0% | 0% | 2.493% | 57
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2024 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1864% | 0.1864% | 5
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2024 | Neither | 13 | 0 | 0 | 0.4845% | 0% | 0% | 0.4845% | 13
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2025 | Neither | 37 | 0 | 0 | 1.173% | 0% | 0% | 1.173% | 37
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Cadiere Forceps | mechanical | 2026 | Neither | 15 | 0 | 0 | 1.142% | 0% | 0% | 1.142% | 15
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2022 | Injury | 0 | 0 | 6 | 0% | 0% | 0.32% | 0.32% | 6
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2022 | Neither | 35 | 0 | 0 | 1.867% | 0% | 0% | 1.867% | 35
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2023 | Injury | 0 | 0 | 11 | 0% | 0% | 0.4812% | 0.4812% | 11
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 51 | 0 | 0 | 2.231% | 0% | 0% | 2.231% | 51
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2024 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2236% | 0.2236% | 6
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 12 | 0 | 0 | 0.4473% | 0% | 0% | 0.4473% | 12
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2025 | Neither | 26 | 0 | 0 | 0.8246% | 0% | 0% | 0.8246% | 26
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 15 | 0 | 0 | 1.142% | 0% | 0% | 1.142% | 15
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2021 | Neither | 45 | 0 | 0 | 2.823% | 0% | 0% | 2.823% | 45
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2022 | Injury | 0 | 0 | 6 | 0% | 0% | 0.32% | 0.32% | 6
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2022 | Neither | 42 | 0 | 0 | 2.24% | 0% | 0% | 2.24% | 42
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2023 | Injury | 0 | 0 | 9 | 0% | 0% | 0.3937% | 0.3937% | 9
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2023 | Neither | 31 | 0 | 0 | 1.356% | 0% | 0% | 1.356% | 31
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2024 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2236% | 0.2236% | 6
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2024 | Neither | 10 | 0 | 0 | 0.3727% | 0% | 0% | 0.3727% | 10
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2025 | Neither | 14 | 0 | 0 | 0.444% | 0% | 0% | 0.444% | 14
Detachment of Device or Device Component | Permanent Cautery Spatula | energy | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2021 | Injury | 0 | 0 | 13 | 0% | 0% | 0.8156% | 0.8156% | 13
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2022 | Injury | 0 | 0 | 11 | 0% | 0% | 0.5867% | 0.5867% | 11
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2023 | Injury | 0 | 0 | 34 | 0% | 0% | 1.487% | 1.487% | 34
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2023 | Neither | 12 | 0 | 0 | 0.5249% | 0% | 0% | 0.5249% | 12
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2024 | Injury | 0 | 0 | 16 | 0% | 0% | 0.5963% | 0.5963% | 16
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2024 | Neither | 13 | 0 | 0 | 0.4845% | 0% | 0% | 0.4845% | 13
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2025 | Injury | 0 | 0 | 23 | 0% | 0% | 0.7295% | 0.7295% | 23
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2025 | Neither | 11 | 0 | 0 | 0.3489% | 0% | 0% | 0.3489% | 11
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2026 | Injury | 0 | 0 | 11 | 0% | 0% | 0.8373% | 0.8373% | 11
Detachment of Device or Device Component | Tip Cover Accessory | energy | 2026 | Neither | 9 | 0 | 0 | 0.6851% | 0% | 0% | 0.6851% | 9
Detachment of Device or Device Component | Vessel Sealer | energy | 2021 | Injury | 0 | 0 | 4 | 0% | 0% | 0.2509% | 0.2509% | 4
Detachment of Device or Device Component | Vessel Sealer | energy | 2022 | Injury | 0 | 0 | 5 | 0% | 0% | 0.2667% | 0.2667% | 5
Detachment of Device or Device Component | Vessel Sealer | energy | 2022 | Neither | 8 | 0 | 0 | 0.4267% | 0% | 0% | 0.4267% | 8
Detachment of Device or Device Component | Vessel Sealer | energy | 2023 | Injury | 0 | 0 | 18 | 0% | 0% | 0.7874% | 0.7874% | 18
Detachment of Device or Device Component | Vessel Sealer | energy | 2023 | Neither | 46 | 0 | 0 | 2.012% | 0% | 0% | 2.012% | 46
Detachment of Device or Device Component | Vessel Sealer | energy | 2024 | Injury | 0 | 0 | 18 | 0% | 0% | 0.6709% | 0.6709% | 18
Detachment of Device or Device Component | Vessel Sealer | energy | 2024 | Neither | 7 | 0 | 0 | 0.2609% | 0% | 0% | 0.2609% | 7
Detachment of Device or Device Component | Vessel Sealer | energy | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Detachment of Device or Device Component | Vessel Sealer | energy | 2025 | Neither | 8 | 0 | 0 | 0.2537% | 0% | 0% | 0.2537% | 8
Detachment of Device or Device Component | Vessel Sealer | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Detachment of Device or Device Component | Vessel Sealer | energy | 2026 | Neither | 9 | 0 | 0 | 0.6851% | 0% | 0% | 0.6851% | 9
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2025 | Neither | 11 | 0 | 0 | 0.3489% | 0% | 0% | 0.3489% | 11
Detachment of Device or Device Component | Patient-Side Cart/Arm | system | 2026 | Neither | 92 | 0 | 0 | 7.003% | 0% | 0% | 7.003% | 92
Detachment of Device or Device Component | SynchroSeal | energy | 2021 | Injury | 0 | 0 | 19 | 0% | 0% | 1.192% | 1.192% | 19
Detachment of Device or Device Component | SynchroSeal | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Detachment of Device or Device Component | SynchroSeal | energy | 2022 | Injury | 0 | 0 | 18 | 0% | 0% | 0.96% | 0.96% | 18
Detachment of Device or Device Component | SynchroSeal | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Detachment of Device or Device Component | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 25 | 0% | 0% | 1.094% | 1.094% | 25
Detachment of Device or Device Component | SynchroSeal | energy | 2023 | Neither | 12 | 0 | 0 | 0.5249% | 0% | 0% | 0.5249% | 12
Detachment of Device or Device Component | SynchroSeal | energy | 2024 | Injury | 0 | 0 | 11 | 0% | 0% | 0.41% | 0.41% | 11
Detachment of Device or Device Component | SynchroSeal | energy | 2024 | Neither | 6 | 0 | 0 | 0.2236% | 0% | 0% | 0.2236% | 6
Detachment of Device or Device Component | SynchroSeal | energy | 2025 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1586% | 0.1586% | 5
Detachment of Device or Device Component | SynchroSeal | energy | 2025 | Neither | 5 | 0 | 0 | 0.1586% | 0% | 0% | 0.1586% | 5
Detachment of Device or Device Component | SynchroSeal | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Detachment of Device or Device Component | SynchroSeal | energy | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Detachment of Device or Device Component | Suction Irrigator | accessory | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Detachment of Device or Device Component | Suction Irrigator | accessory | 2022 | Injury | 0 | 0 | 8 | 0% | 0% | 0.4267% | 0.4267% | 8
Detachment of Device or Device Component | Suction Irrigator | accessory | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Suction Irrigator | accessory | 2023 | Injury | 0 | 0 | 15 | 0% | 0% | 0.6562% | 0.6562% | 15
Detachment of Device or Device Component | Suction Irrigator | accessory | 2023 | Neither | 11 | 0 | 0 | 0.4812% | 0% | 0% | 0.4812% | 11
Detachment of Device or Device Component | Suction Irrigator | accessory | 2024 | Injury | 0 | 0 | 13 | 0% | 0% | 0.4845% | 0.4845% | 13
Detachment of Device or Device Component | Suction Irrigator | accessory | 2024 | Neither | 9 | 0 | 0 | 0.3354% | 0% | 0% | 0.3354% | 9
Detachment of Device or Device Component | Suction Irrigator | accessory | 2025 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1586% | 0.1586% | 5
Detachment of Device or Device Component | Suction Irrigator | accessory | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Detachment of Device or Device Component | Suction Irrigator | accessory | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Suction Irrigator | accessory | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Detachment of Device or Device Component | Bipolar (other) | energy | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Detachment of Device or Device Component | Bipolar (other) | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Detachment of Device or Device Component | Bipolar (other) | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Detachment of Device or Device Component | Bipolar (other) | energy | 2023 | Injury | 0 | 0 | 12 | 0% | 0% | 0.5249% | 0.5249% | 12
Detachment of Device or Device Component | Bipolar (other) | energy | 2023 | Neither | 25 | 0 | 0 | 1.094% | 0% | 0% | 1.094% | 25
Detachment of Device or Device Component | Bipolar (other) | energy | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Detachment of Device or Device Component | Bipolar (other) | energy | 2024 | Neither | 6 | 0 | 0 | 0.2236% | 0% | 0% | 0.2236% | 6
Detachment of Device or Device Component | Bipolar (other) | energy | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Detachment of Device or Device Component | Bipolar (other) | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Bipolar (other) | energy | 2026 | Neither | 7 | 0 | 0 | 0.5328% | 0% | 0% | 0.5328% | 7
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2021 | Injury | 0 | 0 | 6 | 0% | 0% | 0.3764% | 0.3764% | 6
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Injury | 0 | 0 | 7 | 0% | 0% | 0.3733% | 0.3733% | 7
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Injury | 0 | 0 | 10 | 0% | 0% | 0.4374% | 0.4374% | 10
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Neither | 19 | 0 | 0 | 0.8311% | 0% | 0% | 0.8311% | 19
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Neither | 6 | 0 | 0 | 0.2236% | 0% | 0% | 0.2236% | 6
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Neither | 8 | 0 | 0 | 0.2537% | 0% | 0% | 0.2537% | 8
Detachment of Device or Device Component | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Detachment of Device or Device Component | Generic System/Platform | system | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Detachment of Device or Device Component | Generic System/Platform | system | 2021 | Neither | 14 | 0 | 0 | 0.8783% | 0% | 0% | 0.8783% | 14
Detachment of Device or Device Component | Generic System/Platform | system | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Detachment of Device or Device Component | Generic System/Platform | system | 2022 | Neither | 10 | 0 | 0 | 0.5333% | 0% | 0% | 0.5333% | 10
Detachment of Device or Device Component | Generic System/Platform | system | 2023 | Injury | 0 | 0 | 5 | 0% | 0% | 0.2187% | 0.2187% | 5
Detachment of Device or Device Component | Generic System/Platform | system | 2023 | Neither | 6 | 0 | 0 | 0.2625% | 0% | 0% | 0.2625% | 6
Detachment of Device or Device Component | Generic System/Platform | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Detachment of Device or Device Component | Generic System/Platform | system | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Detachment of Device or Device Component | Generic System/Platform | system | 2025 | Neither | 5 | 0 | 0 | 0.1586% | 0% | 0% | 0.1586% | 5
Detachment of Device or Device Component | Generic System/Platform | system | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Detachment of Device or Device Component | Generic System/Platform | system | 2026 | Neither | 7 | 0 | 0 | 0.5328% | 0% | 0% | 0.5328% | 7
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2021 | Injury | 0 | 0 | 10 | 0% | 0% | 0.6274% | 0.6274% | 10
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Injury | 0 | 0 | 8 | 0% | 0% | 0.2982% | 0.2982% | 8
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Detachment of Device or Device Component | Cannula/Trocar/Reducer/Sheath | accessory | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Detachment of Device or Device Component | Endoscope/Camera | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Endoscope/Camera | system | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Detachment of Device or Device Component | Endoscope/Camera | system | 2023 | Neither | 20 | 0 | 0 | 0.8749% | 0% | 0% | 0.8749% | 20
Detachment of Device or Device Component | Endoscope/Camera | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Detachment of Device or Device Component | Endoscope/Camera | system | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Detachment of Device or Device Component | Endoscope/Camera | system | 2025 | Neither | 7 | 0 | 0 | 0.222% | 0% | 0% | 0.222% | 7
Detachment of Device or Device Component | Endoscope/Camera | system | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2022 | Neither | 5 | 0 | 0 | 0.2667% | 0% | 0% | 0.2667% | 5
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2024 | Neither | 11 | 0 | 0 | 0.41% | 0% | 0% | 0.41% | 11
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Detachment of Device or Device Component | Monopolar Cautery (other) | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Detachment of Device or Device Component | Grasper/Forceps (other) | mechanical | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Detachment of Device or Device Component | Curved Bipolar Dissector | energy | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Detachment of Device or Device Component | Cobra Grasper | mechanical | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Detachment of Device or Device Component | Cobra Grasper | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Detachment of Device or Device Component | Cobra Grasper | mechanical | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Detachment of Device or Device Component | Cobra Grasper | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Detachment of Device or Device Component | Cobra Grasper | mechanical | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Detachment of Device or Device Component | PK Dissecting Forceps | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Detachment of Device or Device Component | PK Dissecting Forceps | energy | 2023 | Neither | 6 | 0 | 0 | 0.2625% | 0% | 0% | 0.2625% | 6
Detachment of Device or Device Component | Surgeon Console | system | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Detachment of Device or Device Component | Surgeon Console | system | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Fragmentation | Fenestrated Bipolar | energy | 2021 | Neither | 4 | 0 | 0 | 0.2509% | 0% | 0% | 0.2509% | 4
Material Fragmentation | Fenestrated Bipolar | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Fenestrated Bipolar | energy | 2022 | Neither | 27 | 0 | 0 | 1.44% | 0% | 0% | 1.44% | 27
Material Fragmentation | Fenestrated Bipolar | energy | 2023 | Injury | 0 | 0 | 14 | 0% | 0% | 0.6124% | 0.6124% | 14
Material Fragmentation | Fenestrated Bipolar | energy | 2023 | Neither | 127 | 0 | 0 | 5.556% | 0% | 0% | 5.556% | 127
Material Fragmentation | Fenestrated Bipolar | energy | 2024 | Injury | 0 | 0 | 50 | 0% | 0% | 1.864% | 1.864% | 50
Material Fragmentation | Fenestrated Bipolar | energy | 2024 | Neither | 282 | 0 | 0 | 10.51% | 0% | 0% | 10.51% | 282
Material Fragmentation | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 61 | 0% | 0% | 1.935% | 1.935% | 61
Material Fragmentation | Fenestrated Bipolar | energy | 2025 | Neither | 956 | 0 | 0 | 30.32% | 0% | 0% | 30.32% | 956
Material Fragmentation | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 18 | 0% | 0% | 1.37% | 1.37% | 18
Material Fragmentation | Fenestrated Bipolar | energy | 2026 | Neither | 38 | 0 | 0 | 2.892% | 0% | 0% | 2.892% | 38
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2022 | Neither | 13 | 0 | 0 | 0.6933% | 0% | 0% | 0.6933% | 13
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2023 | Neither | 68 | 0 | 0 | 2.975% | 0% | 0% | 2.975% | 68
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2024 | Injury | 0 | 0 | 25 | 0% | 0% | 0.9318% | 0.9318% | 25
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2024 | Neither | 1115 | 0 | 0 | 41.56% | 0% | 0% | 41.56% | 1115
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 24 | 0% | 0% | 0.7612% | 0.7612% | 24
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2025 | Neither | 266 | 0 | 0 | 8.436% | 0% | 0% | 8.436% | 266
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Material Fragmentation | Needle Driver / SutureCut | mechanical | 2026 | Neither | 13 | 0 | 0 | 0.9895% | 0% | 0% | 0.9895% | 13
Material Fragmentation | Monopolar Curved Scissors | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Monopolar Curved Scissors | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Monopolar Curved Scissors | energy | 2022 | Neither | 12 | 0 | 0 | 0.64% | 0% | 0% | 0.64% | 12
Material Fragmentation | Monopolar Curved Scissors | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Fragmentation | Monopolar Curved Scissors | energy | 2023 | Neither | 29 | 0 | 0 | 1.269% | 0% | 0% | 1.269% | 29
Material Fragmentation | Monopolar Curved Scissors | energy | 2024 | Injury | 0 | 0 | 13 | 0% | 0% | 0.4845% | 0.4845% | 13
Material Fragmentation | Monopolar Curved Scissors | energy | 2024 | Neither | 49 | 0 | 0 | 1.826% | 0% | 0% | 1.826% | 49
Material Fragmentation | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 15 | 0% | 0% | 0.4757% | 0.4757% | 15
Material Fragmentation | Monopolar Curved Scissors | energy | 2025 | Neither | 767 | 0 | 0 | 24.33% | 0% | 0% | 24.33% | 767
Material Fragmentation | Monopolar Curved Scissors | energy | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Material Fragmentation | Monopolar Curved Scissors | energy | 2026 | Neither | 86 | 0 | 0 | 6.546% | 0% | 0% | 6.546% | 86
Material Fragmentation | ProGrasp Forceps | mechanical | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Material Fragmentation | ProGrasp Forceps | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | ProGrasp Forceps | mechanical | 2022 | Injury | 0 | 0 | 5 | 0% | 0% | 0.2667% | 0.2667% | 5
Material Fragmentation | ProGrasp Forceps | mechanical | 2022 | Neither | 9 | 0 | 0 | 0.48% | 0% | 0% | 0.48% | 9
Material Fragmentation | ProGrasp Forceps | mechanical | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Fragmentation | ProGrasp Forceps | mechanical | 2023 | Neither | 12 | 0 | 0 | 0.5249% | 0% | 0% | 0.5249% | 12
Material Fragmentation | ProGrasp Forceps | mechanical | 2024 | Injury | 0 | 0 | 26 | 0% | 0% | 0.9691% | 0.9691% | 26
Material Fragmentation | ProGrasp Forceps | mechanical | 2024 | Neither | 55 | 0 | 0 | 2.05% | 0% | 0% | 2.05% | 55
Material Fragmentation | ProGrasp Forceps | mechanical | 2025 | Injury | 0 | 0 | 25 | 0% | 0% | 0.7929% | 0.7929% | 25
Material Fragmentation | ProGrasp Forceps | mechanical | 2025 | Neither | 555 | 0 | 0 | 17.6% | 0% | 0% | 17.6% | 555
Material Fragmentation | ProGrasp Forceps | mechanical | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Material Fragmentation | ProGrasp Forceps | mechanical | 2026 | Neither | 12 | 0 | 0 | 0.9134% | 0% | 0% | 0.9134% | 12
Material Fragmentation | Maryland Bipolar | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Maryland Bipolar | energy | 2022 | Neither | 18 | 0 | 0 | 0.96% | 0% | 0% | 0.96% | 18
Material Fragmentation | Maryland Bipolar | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Maryland Bipolar | energy | 2023 | Neither | 65 | 0 | 0 | 2.843% | 0% | 0% | 2.843% | 65
Material Fragmentation | Maryland Bipolar | energy | 2024 | Injury | 0 | 0 | 8 | 0% | 0% | 0.2982% | 0.2982% | 8
Material Fragmentation | Maryland Bipolar | energy | 2024 | Neither | 74 | 0 | 0 | 2.758% | 0% | 0% | 2.758% | 74
Material Fragmentation | Maryland Bipolar | energy | 2025 | Injury | 0 | 0 | 14 | 0% | 0% | 0.444% | 0.444% | 14
Material Fragmentation | Maryland Bipolar | energy | 2025 | Neither | 314 | 0 | 0 | 9.959% | 0% | 0% | 9.959% | 314
Material Fragmentation | Maryland Bipolar | energy | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Material Fragmentation | Maryland Bipolar | energy | 2026 | Neither | 17 | 0 | 0 | 1.294% | 0% | 0% | 1.294% | 17
Material Fragmentation | Tenaculum Forceps | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | Tenaculum Forceps | mechanical | 2022 | Injury | 0 | 0 | 4 | 0% | 0% | 0.2133% | 0.2133% | 4
Material Fragmentation | Tenaculum Forceps | mechanical | 2022 | Neither | 15 | 0 | 0 | 0.8% | 0% | 0% | 0.8% | 15
Material Fragmentation | Tenaculum Forceps | mechanical | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Fragmentation | Tenaculum Forceps | mechanical | 2023 | Neither | 70 | 0 | 0 | 3.062% | 0% | 0% | 3.062% | 70
Material Fragmentation | Tenaculum Forceps | mechanical | 2024 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2236% | 0.2236% | 6
Material Fragmentation | Tenaculum Forceps | mechanical | 2024 | Neither | 285 | 0 | 0 | 10.62% | 0% | 0% | 10.62% | 285
Material Fragmentation | Tenaculum Forceps | mechanical | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Fragmentation | Tenaculum Forceps | mechanical | 2025 | Neither | 52 | 0 | 0 | 1.649% | 0% | 0% | 1.649% | 52
Material Fragmentation | Grasping Retractor | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Grasping Retractor | mechanical | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Material Fragmentation | Grasping Retractor | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Grasping Retractor | mechanical | 2022 | Neither | 24 | 0 | 0 | 1.28% | 0% | 0% | 1.28% | 24
Material Fragmentation | Grasping Retractor | mechanical | 2023 | Neither | 53 | 0 | 0 | 2.318% | 0% | 0% | 2.318% | 53
Material Fragmentation | Grasping Retractor | mechanical | 2024 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1864% | 0.1864% | 5
Material Fragmentation | Grasping Retractor | mechanical | 2024 | Neither | 242 | 0 | 0 | 9.02% | 0% | 0% | 9.02% | 242
Material Fragmentation | Grasping Retractor | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Fragmentation | Grasping Retractor | mechanical | 2025 | Neither | 50 | 0 | 0 | 1.586% | 0% | 0% | 1.586% | 50
Material Fragmentation | Force Bipolar | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Force Bipolar | energy | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Material Fragmentation | Force Bipolar | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Force Bipolar | energy | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Material Fragmentation | Force Bipolar | energy | 2023 | Injury | 0 | 0 | 9 | 0% | 0% | 0.3937% | 0.3937% | 9
Material Fragmentation | Force Bipolar | energy | 2023 | Neither | 35 | 0 | 0 | 1.531% | 0% | 0% | 1.531% | 35
Material Fragmentation | Force Bipolar | energy | 2024 | Injury | 0 | 0 | 18 | 0% | 0% | 0.6709% | 0.6709% | 18
Material Fragmentation | Force Bipolar | energy | 2024 | Neither | 44 | 0 | 0 | 1.64% | 0% | 0% | 1.64% | 44
Material Fragmentation | Force Bipolar | energy | 2025 | Injury | 0 | 0 | 36 | 0% | 0% | 1.142% | 1.142% | 36
Material Fragmentation | Force Bipolar | energy | 2025 | Neither | 48 | 0 | 0 | 1.522% | 0% | 0% | 1.522% | 48
Material Fragmentation | Force Bipolar | energy | 2026 | Injury | 0 | 0 | 10 | 0% | 0% | 0.7612% | 0.7612% | 10
Material Fragmentation | Force Bipolar | energy | 2026 | Neither | 6 | 0 | 0 | 0.4567% | 0% | 0% | 0.4567% | 6
Material Fragmentation | Permanent Cautery Hook | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Permanent Cautery Hook | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Permanent Cautery Hook | energy | 2022 | Neither | 11 | 0 | 0 | 0.5867% | 0% | 0% | 0.5867% | 11
Material Fragmentation | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Fragmentation | Permanent Cautery Hook | energy | 2023 | Neither | 43 | 0 | 0 | 1.881% | 0% | 0% | 1.881% | 43
Material Fragmentation | Permanent Cautery Hook | energy | 2024 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3727% | 0.3727% | 10
Material Fragmentation | Permanent Cautery Hook | energy | 2024 | Neither | 59 | 0 | 0 | 2.199% | 0% | 0% | 2.199% | 59
Material Fragmentation | Permanent Cautery Hook | energy | 2025 | Injury | 0 | 0 | 34 | 0% | 0% | 1.078% | 1.078% | 34
Material Fragmentation | Permanent Cautery Hook | energy | 2025 | Neither | 31 | 0 | 0 | 0.9832% | 0% | 0% | 0.9832% | 31
Material Fragmentation | Permanent Cautery Hook | energy | 2026 | Injury | 0 | 0 | 13 | 0% | 0% | 0.9895% | 0.9895% | 13
Material Fragmentation | Permanent Cautery Hook | energy | 2026 | Neither | 9 | 0 | 0 | 0.6851% | 0% | 0% | 0.6851% | 9
Material Fragmentation | Harmonic (ultrasonic) | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Harmonic (ultrasonic) | energy | 2022 | Injury | 0 | 0 | 8 | 0% | 0% | 0.4267% | 0.4267% | 8
Material Fragmentation | Harmonic (ultrasonic) | energy | 2022 | Neither | 5 | 0 | 0 | 0.2667% | 0% | 0% | 0.2667% | 5
Material Fragmentation | Harmonic (ultrasonic) | energy | 2023 | Injury | 0 | 0 | 13 | 0% | 0% | 0.5687% | 0.5687% | 13
Material Fragmentation | Harmonic (ultrasonic) | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Material Fragmentation | Harmonic (ultrasonic) | energy | 2024 | Injury | 0 | 0 | 46 | 0% | 0% | 1.714% | 1.714% | 46
Material Fragmentation | Harmonic (ultrasonic) | energy | 2024 | Neither | 22 | 0 | 0 | 0.82% | 0% | 0% | 0.82% | 22
Material Fragmentation | Harmonic (ultrasonic) | energy | 2025 | Injury | 0 | 0 | 85 | 0% | 0% | 2.696% | 2.696% | 85
Material Fragmentation | Harmonic (ultrasonic) | energy | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Material Fragmentation | Harmonic (ultrasonic) | energy | 2026 | Injury | 0 | 0 | 10 | 0% | 0% | 0.7612% | 0.7612% | 10
Material Fragmentation | Harmonic (ultrasonic) | energy | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Material Fragmentation | Cadiere Forceps | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | Cadiere Forceps | mechanical | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Cadiere Forceps | mechanical | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Material Fragmentation | Cadiere Forceps | mechanical | 2023 | Injury | 0 | 0 | 10 | 0% | 0% | 0.4374% | 0.4374% | 10
Material Fragmentation | Cadiere Forceps | mechanical | 2023 | Neither | 15 | 0 | 0 | 0.6562% | 0% | 0% | 0.6562% | 15
Material Fragmentation | Cadiere Forceps | mechanical | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Material Fragmentation | Cadiere Forceps | mechanical | 2024 | Neither | 9 | 0 | 0 | 0.3354% | 0% | 0% | 0.3354% | 9
Material Fragmentation | Cadiere Forceps | mechanical | 2025 | Injury | 0 | 0 | 7 | 0% | 0% | 0.222% | 0.222% | 7
Material Fragmentation | Cadiere Forceps | mechanical | 2025 | Neither | 81 | 0 | 0 | 2.569% | 0% | 0% | 2.569% | 81
Material Fragmentation | Cadiere Forceps | mechanical | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Material Fragmentation | Cadiere Forceps | mechanical | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2022 | Neither | 9 | 0 | 0 | 0.48% | 0% | 0% | 0.48% | 9
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 13 | 0 | 0 | 0.5687% | 0% | 0% | 0.5687% | 13
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2024 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2236% | 0.2236% | 6
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 14 | 0 | 0 | 0.5218% | 0% | 0% | 0.5218% | 14
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2025 | Injury | 0 | 0 | 13 | 0% | 0% | 0.4123% | 0.4123% | 13
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2025 | Neither | 61 | 0 | 0 | 1.935% | 0% | 0% | 1.935% | 61
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Material Fragmentation | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Material Fragmentation | Clip Applier | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Material Fragmentation | Clip Applier | mechanical | 2022 | Neither | 11 | 0 | 0 | 0.5867% | 0% | 0% | 0.5867% | 11
Material Fragmentation | Clip Applier | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Clip Applier | mechanical | 2023 | Neither | 14 | 0 | 0 | 0.6124% | 0% | 0% | 0.6124% | 14
Material Fragmentation | Clip Applier | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Fragmentation | Clip Applier | mechanical | 2024 | Neither | 28 | 0 | 0 | 1.044% | 0% | 0% | 1.044% | 28
Material Fragmentation | Clip Applier | mechanical | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Fragmentation | Clip Applier | mechanical | 2025 | Neither | 41 | 0 | 0 | 1.3% | 0% | 0% | 1.3% | 41
Material Fragmentation | Clip Applier | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Fragmentation | Clip Applier | mechanical | 2026 | Neither | 5 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 5
Material Fragmentation | Permanent Cautery Spatula | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | Permanent Cautery Spatula | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Permanent Cautery Spatula | energy | 2022 | Neither | 5 | 0 | 0 | 0.2667% | 0% | 0% | 0.2667% | 5
Material Fragmentation | Permanent Cautery Spatula | energy | 2023 | Neither | 20 | 0 | 0 | 0.8749% | 0% | 0% | 0.8749% | 20
Material Fragmentation | Permanent Cautery Spatula | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Fragmentation | Permanent Cautery Spatula | energy | 2024 | Neither | 50 | 0 | 0 | 1.864% | 0% | 0% | 1.864% | 50
Material Fragmentation | Permanent Cautery Spatula | energy | 2025 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3172% | 0.3172% | 10
Material Fragmentation | Permanent Cautery Spatula | energy | 2025 | Neither | 10 | 0 | 0 | 0.3172% | 0% | 0% | 0.3172% | 10
Material Fragmentation | Permanent Cautery Spatula | energy | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Material Fragmentation | Permanent Cautery Spatula | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Fragmentation | Bipolar (other) | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Fragmentation | Bipolar (other) | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Bipolar (other) | energy | 2023 | Neither | 17 | 0 | 0 | 0.7437% | 0% | 0% | 0.7437% | 17
Material Fragmentation | Bipolar (other) | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Fragmentation | Bipolar (other) | energy | 2024 | Neither | 16 | 0 | 0 | 0.5963% | 0% | 0% | 0.5963% | 16
Material Fragmentation | Bipolar (other) | energy | 2025 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3172% | 0.3172% | 10
Material Fragmentation | Bipolar (other) | energy | 2025 | Neither | 46 | 0 | 0 | 1.459% | 0% | 0% | 1.459% | 46
Material Fragmentation | Vessel Sealer | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Vessel Sealer | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Material Fragmentation | Vessel Sealer | energy | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Material Fragmentation | Vessel Sealer | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Material Fragmentation | Vessel Sealer | energy | 2024 | Injury | 0 | 0 | 24 | 0% | 0% | 0.8945% | 0.8945% | 24
Material Fragmentation | Vessel Sealer | energy | 2024 | Neither | 11 | 0 | 0 | 0.41% | 0% | 0% | 0.41% | 11
Material Fragmentation | Vessel Sealer | energy | 2025 | Injury | 0 | 0 | 27 | 0% | 0% | 0.8563% | 0.8563% | 27
Material Fragmentation | Vessel Sealer | energy | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Material Fragmentation | Vessel Sealer | energy | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Material Fragmentation | Vessel Sealer | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Fragmentation | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Material Fragmentation | SynchroSeal | energy | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Material Fragmentation | SynchroSeal | energy | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Material Fragmentation | SynchroSeal | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Fragmentation | SynchroSeal | energy | 2025 | Injury | 0 | 0 | 11 | 0% | 0% | 0.3489% | 0.3489% | 11
Material Fragmentation | SynchroSeal | energy | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Material Fragmentation | SynchroSeal | energy | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Material Fragmentation | SynchroSeal | energy | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Material Fragmentation | Generic System/Platform | system | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Generic System/Platform | system | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Fragmentation | Generic System/Platform | system | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Fragmentation | Generic System/Platform | system | 2024 | Neither | 4 | 0 | 0 | 0.1491% | 0% | 0% | 0.1491% | 4
Material Fragmentation | Generic System/Platform | system | 2025 | Injury | 0 | 0 | 6 | 0% | 0% | 0.1903% | 0.1903% | 6
Material Fragmentation | Generic System/Platform | system | 2025 | Neither | 5 | 0 | 0 | 0.1586% | 0% | 0% | 0.1586% | 5
Material Fragmentation | Generic System/Platform | system | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Fragmentation | Generic System/Platform | system | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Neither | 5 | 0 | 0 | 0.1864% | 0% | 0% | 0.1864% | 5
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1586% | 0.1586% | 5
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Fragmentation | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Material Fragmentation | Curved Bipolar Dissector | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Curved Bipolar Dissector | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Fragmentation | Curved Bipolar Dissector | energy | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Fragmentation | Curved Bipolar Dissector | energy | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Fragmentation | Curved Bipolar Dissector | energy | 2025 | Neither | 16 | 0 | 0 | 0.5075% | 0% | 0% | 0.5075% | 16
Material Fragmentation | Curved Bipolar Dissector | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2024 | Neither | 9 | 0 | 0 | 0.3354% | 0% | 0% | 0.3354% | 9
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2025 | Neither | 5 | 0 | 0 | 0.1586% | 0% | 0% | 0.1586% | 5
Material Fragmentation | Grasper/Forceps (other) | mechanical | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Fragmentation | Suction Irrigator | accessory | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Suction Irrigator | accessory | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Fragmentation | Suction Irrigator | accessory | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Fragmentation | Suction Irrigator | accessory | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Fragmentation | Suction Irrigator | accessory | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Fragmentation | Suction Irrigator | accessory | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Fragmentation | Suction Irrigator | accessory | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Fragmentation | Suction Irrigator | accessory | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Fragmentation | Tip Cover Accessory | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Tip Cover Accessory | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Fragmentation | Tip Cover Accessory | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Fragmentation | Tip Cover Accessory | energy | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Fragmentation | Tip Cover Accessory | energy | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Fragmentation | Tip Cover Accessory | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Fragmentation | Patient-Side Cart/Arm | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Fragmentation | Patient-Side Cart/Arm | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Fragmentation | Patient-Side Cart/Arm | system | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Fragmentation | Patient-Side Cart/Arm | system | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Fragmentation | Patient-Side Cart/Arm | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Fragmentation | Patient-Side Cart/Arm | system | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Fragmentation | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Fragmentation | Patient-Side Cart/Arm | system | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Fragmentation | Cannula/Trocar/Reducer/Sheath | accessory | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Fragmentation | Endoscope/Camera | system | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Fragmentation | Endoscope/Camera | system | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Fragmentation | Endoscope/Camera | system | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Material Fragmentation | Endoscope/Camera | system | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Fragmentation | Endoscope/Camera | system | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Fragmentation | Monopolar Cautery (other) | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Fragmentation | Monopolar Cautery (other) | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Fragmentation | Monopolar Cautery (other) | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Fragmentation | Monopolar Cautery (other) | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Fragmentation | Monopolar Cautery (other) | energy | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Fragmentation | Monopolar Cautery (other) | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Fragmentation | Cobra Grasper | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Fragmentation | Cobra Grasper | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Fragmentation | Cobra Grasper | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Fragmentation | PK Dissecting Forceps | energy | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Break | Harmonic (ultrasonic) | energy | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Break | Harmonic (ultrasonic) | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Break | Harmonic (ultrasonic) | energy | 2022 | Injury | 0 | 0 | 4 | 0% | 0% | 0.2133% | 0.2133% | 4
Break | Harmonic (ultrasonic) | energy | 2022 | Neither | 5 | 0 | 0 | 0.2667% | 0% | 0% | 0.2667% | 5
Break | Harmonic (ultrasonic) | energy | 2023 | Injury | 0 | 0 | 55 | 0% | 0% | 2.406% | 2.406% | 55
Break | Harmonic (ultrasonic) | energy | 2023 | Neither | 173 | 0 | 0 | 7.568% | 0% | 0% | 7.568% | 173
Break | Harmonic (ultrasonic) | energy | 2024 | Injury | 0 | 0 | 32 | 0% | 0% | 1.193% | 1.193% | 32
Break | Harmonic (ultrasonic) | energy | 2024 | Neither | 78 | 0 | 0 | 2.907% | 0% | 0% | 2.907% | 78
Break | Harmonic (ultrasonic) | energy | 2025 | Injury | 0 | 0 | 17 | 0% | 0% | 0.5392% | 0.5392% | 17
Break | Harmonic (ultrasonic) | energy | 2025 | Neither | 242 | 0 | 0 | 7.675% | 0% | 0% | 7.675% | 242
Break | Harmonic (ultrasonic) | energy | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Break | Harmonic (ultrasonic) | energy | 2026 | Neither | 240 | 0 | 0 | 18.27% | 0% | 0% | 18.27% | 240
Break | Fenestrated Bipolar | energy | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Fenestrated Bipolar | energy | 2021 | Neither | 9 | 0 | 0 | 0.5646% | 0% | 0% | 0.5646% | 9
Break | Fenestrated Bipolar | energy | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Break | Fenestrated Bipolar | energy | 2022 | Neither | 7 | 0 | 0 | 0.3733% | 0% | 0% | 0.3733% | 7
Break | Fenestrated Bipolar | energy | 2023 | Injury | 0 | 0 | 7 | 0% | 0% | 0.3062% | 0.3062% | 7
Break | Fenestrated Bipolar | energy | 2023 | Neither | 93 | 0 | 0 | 4.068% | 0% | 0% | 4.068% | 93
Break | Fenestrated Bipolar | energy | 2024 | Injury | 0 | 0 | 20 | 0% | 0% | 0.7454% | 0.7454% | 20
Break | Fenestrated Bipolar | energy | 2024 | Neither | 214 | 0 | 0 | 7.976% | 0% | 0% | 7.976% | 214
Break | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Break | Fenestrated Bipolar | energy | 2025 | Neither | 269 | 0 | 0 | 8.532% | 0% | 0% | 8.532% | 269
Break | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Break | Fenestrated Bipolar | energy | 2026 | Neither | 152 | 0 | 0 | 11.57% | 0% | 0% | 11.57% | 152
Break | Monopolar Curved Scissors | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Break | Monopolar Curved Scissors | energy | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Break | Monopolar Curved Scissors | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Break | Monopolar Curved Scissors | energy | 2023 | Neither | 28 | 0 | 0 | 1.225% | 0% | 0% | 1.225% | 28
Break | Monopolar Curved Scissors | energy | 2024 | Injury | 0 | 0 | 13 | 0% | 0% | 0.4845% | 0.4845% | 13
Break | Monopolar Curved Scissors | energy | 2024 | Neither | 86 | 0 | 0 | 3.205% | 0% | 0% | 3.205% | 86
Break | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | Monopolar Curved Scissors | energy | 2025 | Neither | 222 | 0 | 0 | 7.041% | 0% | 0% | 7.041% | 222
Break | Monopolar Curved Scissors | energy | 2026 | Neither | 158 | 0 | 0 | 12.03% | 0% | 0% | 12.03% | 158
Break | Force Bipolar | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Break | Force Bipolar | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Break | Force Bipolar | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Break | Force Bipolar | energy | 2022 | Neither | 7 | 0 | 0 | 0.3733% | 0% | 0% | 0.3733% | 7
Break | Force Bipolar | energy | 2023 | Injury | 0 | 0 | 8 | 0% | 0% | 0.35% | 0.35% | 8
Break | Force Bipolar | energy | 2023 | Neither | 78 | 0 | 0 | 3.412% | 0% | 0% | 3.412% | 78
Break | Force Bipolar | energy | 2024 | Injury | 0 | 0 | 8 | 0% | 0% | 0.2982% | 0.2982% | 8
Break | Force Bipolar | energy | 2024 | Neither | 102 | 0 | 0 | 3.802% | 0% | 0% | 3.802% | 102
Break | Force Bipolar | energy | 2025 | Injury | 0 | 0 | 9 | 0% | 0% | 0.2854% | 0.2854% | 9
Break | Force Bipolar | energy | 2025 | Neither | 171 | 0 | 0 | 5.423% | 0% | 0% | 5.423% | 171
Break | Force Bipolar | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Force Bipolar | energy | 2026 | Neither | 89 | 0 | 0 | 6.775% | 0% | 0% | 6.775% | 89
Break | Maryland Bipolar | energy | 2021 | Neither | 4 | 0 | 0 | 0.2509% | 0% | 0% | 0.2509% | 4
Break | Maryland Bipolar | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Break | Maryland Bipolar | energy | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Break | Maryland Bipolar | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Maryland Bipolar | energy | 2023 | Neither | 49 | 0 | 0 | 2.143% | 0% | 0% | 2.143% | 49
Break | Maryland Bipolar | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Maryland Bipolar | energy | 2024 | Neither | 73 | 0 | 0 | 2.721% | 0% | 0% | 2.721% | 73
Break | Maryland Bipolar | energy | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Break | Maryland Bipolar | energy | 2025 | Neither | 135 | 0 | 0 | 4.282% | 0% | 0% | 4.282% | 135
Break | Maryland Bipolar | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Maryland Bipolar | energy | 2026 | Neither | 51 | 0 | 0 | 3.882% | 0% | 0% | 3.882% | 51
Break | Generic System/Platform | system | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Generic System/Platform | system | 2021 | Neither | 71 | 0 | 0 | 4.454% | 0% | 0% | 4.454% | 71
Break | Generic System/Platform | system | 2022 | Injury | 0 | 0 | 5 | 0% | 0% | 0.2667% | 0.2667% | 5
Break | Generic System/Platform | system | 2022 | Neither | 49 | 0 | 0 | 2.613% | 0% | 0% | 2.613% | 49
Break | Generic System/Platform | system | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Break | Generic System/Platform | system | 2023 | Neither | 36 | 0 | 0 | 1.575% | 0% | 0% | 1.575% | 36
Break | Generic System/Platform | system | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Generic System/Platform | system | 2024 | Neither | 6 | 0 | 0 | 0.2236% | 0% | 0% | 0.2236% | 6
Break | Generic System/Platform | system | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Break | Generic System/Platform | system | 2025 | Neither | 94 | 0 | 0 | 2.981% | 0% | 0% | 2.981% | 94
Break | Generic System/Platform | system | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Break | Generic System/Platform | system | 2026 | Neither | 35 | 0 | 0 | 2.664% | 0% | 0% | 2.664% | 35
Break | Permanent Cautery Hook | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Break | Permanent Cautery Hook | energy | 2021 | Neither | 12 | 0 | 0 | 0.7528% | 0% | 0% | 0.7528% | 12
Break | Permanent Cautery Hook | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Break | Permanent Cautery Hook | energy | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Break | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Break | Permanent Cautery Hook | energy | 2023 | Neither | 65 | 0 | 0 | 2.843% | 0% | 0% | 2.843% | 65
Break | Permanent Cautery Hook | energy | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Break | Permanent Cautery Hook | energy | 2024 | Neither | 48 | 0 | 0 | 1.789% | 0% | 0% | 1.789% | 48
Break | Permanent Cautery Hook | energy | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Break | Permanent Cautery Hook | energy | 2025 | Neither | 103 | 0 | 0 | 3.267% | 0% | 0% | 3.267% | 103
Break | Permanent Cautery Hook | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Permanent Cautery Hook | energy | 2026 | Neither | 47 | 0 | 0 | 3.578% | 0% | 0% | 3.578% | 47
Break | Needle Driver / SutureCut | mechanical | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Needle Driver / SutureCut | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Break | Needle Driver / SutureCut | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Break | Needle Driver / SutureCut | mechanical | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Break | Needle Driver / SutureCut | mechanical | 2023 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2625% | 0.2625% | 6
Break | Needle Driver / SutureCut | mechanical | 2023 | Neither | 37 | 0 | 0 | 1.619% | 0% | 0% | 1.619% | 37
Break | Needle Driver / SutureCut | mechanical | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Break | Needle Driver / SutureCut | mechanical | 2024 | Neither | 57 | 0 | 0 | 2.124% | 0% | 0% | 2.124% | 57
Break | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Break | Needle Driver / SutureCut | mechanical | 2025 | Neither | 102 | 0 | 0 | 3.235% | 0% | 0% | 3.235% | 102
Break | Needle Driver / SutureCut | mechanical | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Break | Needle Driver / SutureCut | mechanical | 2026 | Neither | 48 | 0 | 0 | 3.654% | 0% | 0% | 3.654% | 48
Break | ProGrasp Forceps | mechanical | 2022 | Neither | 7 | 0 | 0 | 0.3733% | 0% | 0% | 0.3733% | 7
Break | ProGrasp Forceps | mechanical | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Break | ProGrasp Forceps | mechanical | 2023 | Neither | 34 | 0 | 0 | 1.487% | 0% | 0% | 1.487% | 34
Break | ProGrasp Forceps | mechanical | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Break | ProGrasp Forceps | mechanical | 2024 | Neither | 31 | 0 | 0 | 1.155% | 0% | 0% | 1.155% | 31
Break | ProGrasp Forceps | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | ProGrasp Forceps | mechanical | 2025 | Neither | 96 | 0 | 0 | 3.045% | 0% | 0% | 3.045% | 96
Break | ProGrasp Forceps | mechanical | 2026 | Neither | 52 | 0 | 0 | 3.958% | 0% | 0% | 3.958% | 52
Break | Cadiere Forceps | mechanical | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Cadiere Forceps | mechanical | 2021 | Neither | 3 | 0 | 0 | 0.1882% | 0% | 0% | 0.1882% | 3
Break | Cadiere Forceps | mechanical | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Break | Cadiere Forceps | mechanical | 2023 | Injury | 0 | 0 | 8 | 0% | 0% | 0.35% | 0.35% | 8
Break | Cadiere Forceps | mechanical | 2023 | Neither | 52 | 0 | 0 | 2.275% | 0% | 0% | 2.275% | 52
Break | Cadiere Forceps | mechanical | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Cadiere Forceps | mechanical | 2024 | Neither | 14 | 0 | 0 | 0.5218% | 0% | 0% | 0.5218% | 14
Break | Cadiere Forceps | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Break | Cadiere Forceps | mechanical | 2025 | Neither | 38 | 0 | 0 | 1.205% | 0% | 0% | 1.205% | 38
Break | Cadiere Forceps | mechanical | 2026 | Neither | 28 | 0 | 0 | 2.131% | 0% | 0% | 2.131% | 28
Break | Tip-Up Fenestrated Grasper | mechanical | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Tip-Up Fenestrated Grasper | mechanical | 2022 | Neither | 9 | 0 | 0 | 0.48% | 0% | 0% | 0.48% | 9
Break | Tip-Up Fenestrated Grasper | mechanical | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Break | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 21 | 0 | 0 | 0.9186% | 0% | 0% | 0.9186% | 21
Break | Tip-Up Fenestrated Grasper | mechanical | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Break | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 14 | 0 | 0 | 0.5218% | 0% | 0% | 0.5218% | 14
Break | Tip-Up Fenestrated Grasper | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | Tip-Up Fenestrated Grasper | mechanical | 2025 | Neither | 48 | 0 | 0 | 1.522% | 0% | 0% | 1.522% | 48
Break | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 31 | 0 | 0 | 2.36% | 0% | 0% | 2.36% | 31
Break | Monopolar Cautery (other) | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Break | Monopolar Cautery (other) | energy | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Break | Monopolar Cautery (other) | energy | 2024 | Neither | 20 | 0 | 0 | 0.7454% | 0% | 0% | 0.7454% | 20
Break | Monopolar Cautery (other) | energy | 2025 | Neither | 54 | 0 | 0 | 1.713% | 0% | 0% | 1.713% | 54
Break | Monopolar Cautery (other) | energy | 2026 | Neither | 21 | 0 | 0 | 1.598% | 0% | 0% | 1.598% | 21
Break | Clip Applier | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Clip Applier | mechanical | 2023 | Neither | 12 | 0 | 0 | 0.5249% | 0% | 0% | 0.5249% | 12
Break | Clip Applier | mechanical | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Clip Applier | mechanical | 2024 | Neither | 12 | 0 | 0 | 0.4473% | 0% | 0% | 0.4473% | 12
Break | Clip Applier | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | Clip Applier | mechanical | 2025 | Neither | 40 | 0 | 0 | 1.269% | 0% | 0% | 1.269% | 40
Break | Clip Applier | mechanical | 2026 | Neither | 27 | 0 | 0 | 2.055% | 0% | 0% | 2.055% | 27
Break | Permanent Cautery Spatula | energy | 2021 | Neither | 4 | 0 | 0 | 0.2509% | 0% | 0% | 0.2509% | 4
Break | Permanent Cautery Spatula | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Break | Permanent Cautery Spatula | energy | 2023 | Neither | 29 | 0 | 0 | 1.269% | 0% | 0% | 1.269% | 29
Break | Permanent Cautery Spatula | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Permanent Cautery Spatula | energy | 2024 | Neither | 11 | 0 | 0 | 0.41% | 0% | 0% | 0.41% | 11
Break | Permanent Cautery Spatula | energy | 2025 | Neither | 28 | 0 | 0 | 0.888% | 0% | 0% | 0.888% | 28
Break | Permanent Cautery Spatula | energy | 2026 | Neither | 10 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 10
Break | Bipolar (other) | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Break | Bipolar (other) | energy | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Break | Bipolar (other) | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Bipolar (other) | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Bipolar (other) | energy | 2023 | Neither | 26 | 0 | 0 | 1.137% | 0% | 0% | 1.137% | 26
Break | Bipolar (other) | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Break | Bipolar (other) | energy | 2024 | Neither | 16 | 0 | 0 | 0.5963% | 0% | 0% | 0.5963% | 16
Break | Bipolar (other) | energy | 2025 | Neither | 28 | 0 | 0 | 0.888% | 0% | 0% | 0.888% | 28
Break | Bipolar (other) | energy | 2026 | Neither | 8 | 0 | 0 | 0.6089% | 0% | 0% | 0.6089% | 8
Break | Vessel Sealer | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Break | Vessel Sealer | energy | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Break | Vessel Sealer | energy | 2023 | Neither | 33 | 0 | 0 | 1.444% | 0% | 0% | 1.444% | 33
Break | Vessel Sealer | energy | 2024 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1864% | 0.1864% | 5
Break | Vessel Sealer | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Break | Vessel Sealer | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | Vessel Sealer | energy | 2025 | Neither | 13 | 0 | 0 | 0.4123% | 0% | 0% | 0.4123% | 13
Break | Vessel Sealer | energy | 2026 | Neither | 6 | 0 | 0 | 0.4567% | 0% | 0% | 0.4567% | 6
Break | Stapler (SureForm/EndoWrist) | mechanical | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Break | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Break | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Neither | 12 | 0 | 0 | 0.5249% | 0% | 0% | 0.5249% | 12
Break | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Break | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Neither | 7 | 0 | 0 | 0.2609% | 0% | 0% | 0.2609% | 7
Break | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Neither | 21 | 0 | 0 | 0.666% | 0% | 0% | 0.666% | 21
Break | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Neither | 13 | 0 | 0 | 0.9895% | 0% | 0% | 0.9895% | 13
Break | Tenaculum Forceps | mechanical | 2021 | Neither | 6 | 0 | 0 | 0.3764% | 0% | 0% | 0.3764% | 6
Break | Tenaculum Forceps | mechanical | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Break | Tenaculum Forceps | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Tenaculum Forceps | mechanical | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Break | Tenaculum Forceps | mechanical | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Tenaculum Forceps | mechanical | 2024 | Neither | 15 | 0 | 0 | 0.5591% | 0% | 0% | 0.5591% | 15
Break | Tenaculum Forceps | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Break | Tenaculum Forceps | mechanical | 2025 | Neither | 8 | 0 | 0 | 0.2537% | 0% | 0% | 0.2537% | 8
Break | Tenaculum Forceps | mechanical | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Break | Suction Irrigator | accessory | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Suction Irrigator | accessory | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Break | Suction Irrigator | accessory | 2024 | Neither | 10 | 0 | 0 | 0.3727% | 0% | 0% | 0.3727% | 10
Break | Suction Irrigator | accessory | 2025 | Neither | 19 | 0 | 0 | 0.6026% | 0% | 0% | 0.6026% | 19
Break | Suction Irrigator | accessory | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Suction Irrigator | accessory | 2026 | Neither | 5 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 5
Break | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | SynchroSeal | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Break | SynchroSeal | energy | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Break | SynchroSeal | energy | 2024 | Neither | 8 | 0 | 0 | 0.2982% | 0% | 0% | 0.2982% | 8
Break | SynchroSeal | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | SynchroSeal | energy | 2025 | Neither | 20 | 0 | 0 | 0.6343% | 0% | 0% | 0.6343% | 20
Break | SynchroSeal | energy | 2026 | Neither | 7 | 0 | 0 | 0.5328% | 0% | 0% | 0.5328% | 7
Break | Grasping Retractor | mechanical | 2021 | Neither | 8 | 0 | 0 | 0.5019% | 0% | 0% | 0.5019% | 8
Break | Grasping Retractor | mechanical | 2023 | Neither | 6 | 0 | 0 | 0.2625% | 0% | 0% | 0.2625% | 6
Break | Grasping Retractor | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Break | Grasping Retractor | mechanical | 2024 | Neither | 11 | 0 | 0 | 0.41% | 0% | 0% | 0.41% | 11
Break | Grasping Retractor | mechanical | 2025 | Neither | 9 | 0 | 0 | 0.2854% | 0% | 0% | 0.2854% | 9
Break | Tip Cover Accessory | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Tip Cover Accessory | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Break | Tip Cover Accessory | energy | 2024 | Neither | 10 | 0 | 0 | 0.3727% | 0% | 0% | 0.3727% | 10
Break | Tip Cover Accessory | energy | 2025 | Neither | 9 | 0 | 0 | 0.2854% | 0% | 0% | 0.2854% | 9
Break | Tip Cover Accessory | energy | 2026 | Neither | 9 | 0 | 0 | 0.6851% | 0% | 0% | 0.6851% | 9
Break | Patient-Side Cart/Arm | system | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Break | Patient-Side Cart/Arm | system | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Patient-Side Cart/Arm | system | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Break | Patient-Side Cart/Arm | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Break | Patient-Side Cart/Arm | system | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Break | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Break | Patient-Side Cart/Arm | system | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Patient-Side Cart/Arm | system | 2026 | Neither | 8 | 0 | 0 | 0.6089% | 0% | 0% | 0.6089% | 8
Break | Endoscope/Camera | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Endoscope/Camera | system | 2023 | Neither | 8 | 0 | 0 | 0.35% | 0% | 0% | 0.35% | 8
Break | Endoscope/Camera | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Break | Endoscope/Camera | system | 2025 | Neither | 5 | 0 | 0 | 0.1586% | 0% | 0% | 0.1586% | 5
Break | Endoscope/Camera | system | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Break | Curved Bipolar Dissector | energy | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Break | Curved Bipolar Dissector | energy | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Break | Curved Bipolar Dissector | energy | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Break | Curved Bipolar Dissector | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Curved Bipolar Dissector | energy | 2026 | Neither | 5 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 5
Break | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Break | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Break | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Break | Cannula/Trocar/Reducer/Sheath | accessory | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Break | Cobra Grasper | mechanical | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Break | Cobra Grasper | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Cobra Grasper | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Break | Cobra Grasper | mechanical | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Break | Cobra Grasper | mechanical | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Break | PK Dissecting Forceps | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Break | PK Dissecting Forceps | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Break | Surgeon Console | system | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Break | Surgeon Console | system | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Break | Vision System | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Break | Vision System | system | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Vision System | system | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Break | Vision System | system | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Break | Grasper/Forceps (other) | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Break | Grasper/Forceps (other) | mechanical | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Break | Grasper/Forceps (other) | mechanical | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Break | Unknown/Unspecified | unknown | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Material Separation | Tip Cover Accessory | energy | 2021 | Injury | 0 | 0 | 23 | 0% | 0% | 1.443% | 1.443% | 23
Material Separation | Tip Cover Accessory | energy | 2022 | Injury | 0 | 0 | 21 | 0% | 0% | 1.12% | 1.12% | 21
Material Separation | Tip Cover Accessory | energy | 2023 | Injury | 0 | 0 | 13 | 0% | 0% | 0.5687% | 0.5687% | 13
Material Separation | Tip Cover Accessory | energy | 2023 | Neither | 8 | 0 | 0 | 0.35% | 0% | 0% | 0.35% | 8
Material Separation | Tip Cover Accessory | energy | 2024 | Injury | 0 | 0 | 11 | 0% | 0% | 0.41% | 0.41% | 11
Material Separation | Tip Cover Accessory | energy | 2024 | Neither | 4 | 0 | 0 | 0.1491% | 0% | 0% | 0.1491% | 4
Material Separation | Tip Cover Accessory | energy | 2025 | Injury | 0 | 0 | 55 | 0% | 0% | 1.744% | 1.744% | 55
Material Separation | Tip Cover Accessory | energy | 2025 | Neither | 11 | 0 | 0 | 0.3489% | 0% | 0% | 0.3489% | 11
Material Separation | Tip Cover Accessory | energy | 2026 | Injury | 0 | 0 | 24 | 0% | 0% | 1.827% | 1.827% | 24
Material Separation | Tip Cover Accessory | energy | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Material Separation | Harmonic (ultrasonic) | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Harmonic (ultrasonic) | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Material Separation | Harmonic (ultrasonic) | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Separation | Harmonic (ultrasonic) | energy | 2023 | Neither | 8 | 0 | 0 | 0.35% | 0% | 0% | 0.35% | 8
Material Separation | Harmonic (ultrasonic) | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Separation | Harmonic (ultrasonic) | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Separation | Harmonic (ultrasonic) | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Harmonic (ultrasonic) | energy | 2026 | Injury | 0 | 0 | 13 | 0% | 0% | 0.9895% | 0.9895% | 13
Material Separation | Harmonic (ultrasonic) | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Monopolar Curved Scissors | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Separation | Monopolar Curved Scissors | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Monopolar Curved Scissors | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Separation | Monopolar Curved Scissors | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Monopolar Curved Scissors | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Separation | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 11 | 0% | 0% | 0.3489% | 0.3489% | 11
Material Separation | Monopolar Curved Scissors | energy | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Separation | Monopolar Curved Scissors | energy | 2026 | Injury | 0 | 0 | 6 | 0% | 0% | 0.4567% | 0.4567% | 6
Material Separation | Monopolar Curved Scissors | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Fenestrated Bipolar | energy | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Material Separation | Fenestrated Bipolar | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Separation | Fenestrated Bipolar | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Fenestrated Bipolar | energy | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Material Separation | Fenestrated Bipolar | energy | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Material Separation | Fenestrated Bipolar | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Separation | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Material Separation | Clip Applier | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Clip Applier | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Clip Applier | mechanical | 2023 | Injury | 0 | 0 | 6 | 0% | 0% | 0.2625% | 0.2625% | 6
Material Separation | Clip Applier | mechanical | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Clip Applier | mechanical | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Material Separation | Clip Applier | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Clip Applier | mechanical | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Separation | Clip Applier | mechanical | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Material Separation | Force Bipolar | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Force Bipolar | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Force Bipolar | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Separation | Force Bipolar | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Force Bipolar | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Separation | Force Bipolar | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Separation | Force Bipolar | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Force Bipolar | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Material Separation | Permanent Cautery Hook | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Separation | Permanent Cautery Hook | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Material Separation | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Permanent Cautery Hook | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Permanent Cautery Hook | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Permanent Cautery Hook | energy | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Separation | Permanent Cautery Hook | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Material Separation | Needle Driver / SutureCut | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Separation | Needle Driver / SutureCut | mechanical | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Separation | Needle Driver / SutureCut | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Needle Driver / SutureCut | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Needle Driver / SutureCut | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Separation | Needle Driver / SutureCut | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Separation | Needle Driver / SutureCut | mechanical | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Separation | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Material Separation | Needle Driver / SutureCut | mechanical | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | SynchroSeal | energy | 2021 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1255% | 0.1255% | 2
Material Separation | SynchroSeal | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | SynchroSeal | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | SynchroSeal | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Separation | SynchroSeal | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | SynchroSeal | energy | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Separation | SynchroSeal | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Separation | ProGrasp Forceps | mechanical | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Separation | ProGrasp Forceps | mechanical | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Material Separation | ProGrasp Forceps | mechanical | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Material Separation | ProGrasp Forceps | mechanical | 2024 | Neither | 4 | 0 | 0 | 0.1491% | 0% | 0% | 0.1491% | 4
Material Separation | ProGrasp Forceps | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Separation | Generic System/Platform | system | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Material Separation | Generic System/Platform | system | 2021 | Neither | 3 | 0 | 0 | 0.1882% | 0% | 0% | 0.1882% | 3
Material Separation | Generic System/Platform | system | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Generic System/Platform | system | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Generic System/Platform | system | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Separation | Generic System/Platform | system | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Separation | Generic System/Platform | system | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Patient-Side Cart/Arm | system | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Material Separation | Patient-Side Cart/Arm | system | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Patient-Side Cart/Arm | system | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Separation | Patient-Side Cart/Arm | system | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Separation | Patient-Side Cart/Arm | system | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Cadiere Forceps | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Cadiere Forceps | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Cadiere Forceps | mechanical | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Material Separation | Cadiere Forceps | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Cadiere Forceps | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Cadiere Forceps | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Separation | Maryland Bipolar | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Separation | Maryland Bipolar | energy | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Material Separation | Maryland Bipolar | energy | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Separation | Bipolar (other) | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Separation | Bipolar (other) | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Bipolar (other) | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Separation | Bipolar (other) | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Bipolar (other) | energy | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Separation | Permanent Cautery Spatula | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Separation | Permanent Cautery Spatula | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Permanent Cautery Spatula | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Separation | Vessel Sealer | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Separation | Vessel Sealer | energy | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Vessel Sealer | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Vessel Sealer | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Separation | Surgeon Console | system | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Separation | Surgeon Console | system | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Separation | Surgeon Console | system | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Tip-Up Fenestrated Grasper | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Separation | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Separation | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Tip-Up Fenestrated Grasper | mechanical | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Separation | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Endoscope/Camera | system | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Material Separation | Endoscope/Camera | system | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Separation | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Cannula/Trocar/Reducer/Sheath | accessory | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Separation | Monopolar Cautery (other) | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Separation | Monopolar Cautery (other) | energy | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Separation | Suction Irrigator | accessory | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Separation | Suction Irrigator | accessory | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Separation | Vision System | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Separation | Vision System | system | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Separation | Tenaculum Forceps | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Separation | Grasping Retractor | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Unintended System Motion | Patient-Side Cart/Arm | system | 2021 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1882% | 0.1882% | 3
Unintended System Motion | Patient-Side Cart/Arm | system | 2022 | Injury | 0 | 0 | 7 | 0% | 0% | 0.3733% | 0.3733% | 7
Unintended System Motion | Patient-Side Cart/Arm | system | 2022 | Neither | 82 | 0 | 0 | 4.373% | 0% | 0% | 4.373% | 82
Unintended System Motion | Patient-Side Cart/Arm | system | 2023 | Death | 0 | 1 | 0 | 0% | 0.04374% | 0% | 0.04374% | 1
Unintended System Motion | Patient-Side Cart/Arm | system | 2023 | Injury | 0 | 0 | 10 | 0% | 0% | 0.4374% | 0.4374% | 10
Unintended System Motion | Patient-Side Cart/Arm | system | 2023 | Neither | 131 | 0 | 0 | 5.731% | 0% | 0% | 5.731% | 131
Unintended System Motion | Patient-Side Cart/Arm | system | 2024 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3727% | 0.3727% | 10
Unintended System Motion | Patient-Side Cart/Arm | system | 2024 | Neither | 130 | 0 | 0 | 4.845% | 0% | 0% | 4.845% | 130
Unintended System Motion | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 11 | 0% | 0% | 0.3489% | 0.3489% | 11
Unintended System Motion | Patient-Side Cart/Arm | system | 2025 | Neither | 179 | 0 | 0 | 5.677% | 0% | 0% | 5.677% | 179
Unintended System Motion | Patient-Side Cart/Arm | system | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Unintended System Motion | Patient-Side Cart/Arm | system | 2026 | Neither | 124 | 0 | 0 | 9.439% | 0% | 0% | 9.439% | 124
Unintended System Motion | Surgeon Console | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Surgeon Console | system | 2022 | Neither | 19 | 0 | 0 | 1.013% | 0% | 0% | 1.013% | 19
Unintended System Motion | Surgeon Console | system | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Unintended System Motion | Surgeon Console | system | 2023 | Neither | 42 | 0 | 0 | 1.837% | 0% | 0% | 1.837% | 42
Unintended System Motion | Surgeon Console | system | 2024 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1118% | 0.1118% | 3
Unintended System Motion | Surgeon Console | system | 2024 | Neither | 39 | 0 | 0 | 1.454% | 0% | 0% | 1.454% | 39
Unintended System Motion | Surgeon Console | system | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Surgeon Console | system | 2025 | Neither | 53 | 0 | 0 | 1.681% | 0% | 0% | 1.681% | 53
Unintended System Motion | Surgeon Console | system | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Unintended System Motion | Surgeon Console | system | 2026 | Neither | 48 | 0 | 0 | 3.654% | 0% | 0% | 3.654% | 48
Unintended System Motion | Fenestrated Bipolar | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Fenestrated Bipolar | energy | 2022 | Neither | 13 | 0 | 0 | 0.6933% | 0% | 0% | 0.6933% | 13
Unintended System Motion | Fenestrated Bipolar | energy | 2023 | Neither | 39 | 0 | 0 | 1.706% | 0% | 0% | 1.706% | 39
Unintended System Motion | Fenestrated Bipolar | energy | 2024 | Neither | 36 | 0 | 0 | 1.342% | 0% | 0% | 1.342% | 36
Unintended System Motion | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Fenestrated Bipolar | energy | 2025 | Neither | 70 | 0 | 0 | 2.22% | 0% | 0% | 2.22% | 70
Unintended System Motion | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Unintended System Motion | Fenestrated Bipolar | energy | 2026 | Neither | 31 | 0 | 0 | 2.36% | 0% | 0% | 2.36% | 31
Unintended System Motion | Clip Applier | mechanical | 2021 | Neither | 20 | 0 | 0 | 1.255% | 0% | 0% | 1.255% | 20
Unintended System Motion | Clip Applier | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Unintended System Motion | Clip Applier | mechanical | 2022 | Neither | 45 | 0 | 0 | 2.4% | 0% | 0% | 2.4% | 45
Unintended System Motion | Clip Applier | mechanical | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Unintended System Motion | Clip Applier | mechanical | 2023 | Neither | 38 | 0 | 0 | 1.662% | 0% | 0% | 1.662% | 38
Unintended System Motion | Clip Applier | mechanical | 2024 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1491% | 0.1491% | 4
Unintended System Motion | Clip Applier | mechanical | 2024 | Neither | 40 | 0 | 0 | 1.491% | 0% | 0% | 1.491% | 40
Unintended System Motion | Clip Applier | mechanical | 2025 | Neither | 25 | 0 | 0 | 0.7929% | 0% | 0% | 0.7929% | 25
Unintended System Motion | Clip Applier | mechanical | 2026 | Neither | 13 | 0 | 0 | 0.9895% | 0% | 0% | 0.9895% | 13
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2022 | Neither | 12 | 0 | 0 | 0.64% | 0% | 0% | 0.64% | 12
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2023 | Neither | 83 | 0 | 0 | 3.631% | 0% | 0% | 3.631% | 83
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2024 | Neither | 38 | 0 | 0 | 1.416% | 0% | 0% | 1.416% | 38
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2025 | Neither | 32 | 0 | 0 | 1.015% | 0% | 0% | 1.015% | 32
Unintended System Motion | Needle Driver / SutureCut | mechanical | 2026 | Neither | 22 | 0 | 0 | 1.675% | 0% | 0% | 1.675% | 22
Unintended System Motion | Vessel Sealer | energy | 2022 | Neither | 12 | 0 | 0 | 0.64% | 0% | 0% | 0.64% | 12
Unintended System Motion | Vessel Sealer | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Unintended System Motion | Vessel Sealer | energy | 2023 | Neither | 30 | 0 | 0 | 1.312% | 0% | 0% | 1.312% | 30
Unintended System Motion | Vessel Sealer | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Unintended System Motion | Vessel Sealer | energy | 2024 | Neither | 30 | 0 | 0 | 1.118% | 0% | 0% | 1.118% | 30
Unintended System Motion | Vessel Sealer | energy | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Unintended System Motion | Vessel Sealer | energy | 2025 | Neither | 69 | 0 | 0 | 2.188% | 0% | 0% | 2.188% | 69
Unintended System Motion | Vessel Sealer | energy | 2026 | Neither | 28 | 0 | 0 | 2.131% | 0% | 0% | 2.131% | 28
Unintended System Motion | Monopolar Curved Scissors | energy | 2022 | Neither | 14 | 0 | 0 | 0.7467% | 0% | 0% | 0.7467% | 14
Unintended System Motion | Monopolar Curved Scissors | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Unintended System Motion | Monopolar Curved Scissors | energy | 2023 | Neither | 38 | 0 | 0 | 1.662% | 0% | 0% | 1.662% | 38
Unintended System Motion | Monopolar Curved Scissors | energy | 2024 | Neither | 35 | 0 | 0 | 1.305% | 0% | 0% | 1.305% | 35
Unintended System Motion | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Monopolar Curved Scissors | energy | 2025 | Neither | 36 | 0 | 0 | 1.142% | 0% | 0% | 1.142% | 36
Unintended System Motion | Monopolar Curved Scissors | energy | 2026 | Neither | 19 | 0 | 0 | 1.446% | 0% | 0% | 1.446% | 19
Unintended System Motion | Maryland Bipolar | energy | 2022 | Neither | 10 | 0 | 0 | 0.5333% | 0% | 0% | 0.5333% | 10
Unintended System Motion | Maryland Bipolar | energy | 2023 | Neither | 36 | 0 | 0 | 1.575% | 0% | 0% | 1.575% | 36
Unintended System Motion | Maryland Bipolar | energy | 2024 | Neither | 17 | 0 | 0 | 0.6336% | 0% | 0% | 0.6336% | 17
Unintended System Motion | Maryland Bipolar | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Maryland Bipolar | energy | 2025 | Neither | 35 | 0 | 0 | 1.11% | 0% | 0% | 1.11% | 35
Unintended System Motion | Maryland Bipolar | energy | 2026 | Neither | 22 | 0 | 0 | 1.675% | 0% | 0% | 1.675% | 22
Unintended System Motion | ProGrasp Forceps | mechanical | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Unintended System Motion | ProGrasp Forceps | mechanical | 2023 | Neither | 31 | 0 | 0 | 1.356% | 0% | 0% | 1.356% | 31
Unintended System Motion | ProGrasp Forceps | mechanical | 2024 | Neither | 17 | 0 | 0 | 0.6336% | 0% | 0% | 0.6336% | 17
Unintended System Motion | ProGrasp Forceps | mechanical | 2025 | Neither | 44 | 0 | 0 | 1.395% | 0% | 0% | 1.395% | 44
Unintended System Motion | ProGrasp Forceps | mechanical | 2026 | Neither | 20 | 0 | 0 | 1.522% | 0% | 0% | 1.522% | 20
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Injury | 0 | 0 | 3 | 0% | 0% | 0.16% | 0.16% | 3
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 35 | 0 | 0 | 1.867% | 0% | 0% | 1.867% | 35
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Neither | 33 | 0 | 0 | 1.444% | 0% | 0% | 1.444% | 33
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Neither | 17 | 0 | 0 | 0.6336% | 0% | 0% | 0.6336% | 17
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Neither | 12 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 12
Unintended System Motion | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Neither | 10 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 10
Unintended System Motion | Endoscope/Camera | system | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Unintended System Motion | Endoscope/Camera | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Endoscope/Camera | system | 2022 | Neither | 12 | 0 | 0 | 0.64% | 0% | 0% | 0.64% | 12
Unintended System Motion | Endoscope/Camera | system | 2023 | Neither | 43 | 0 | 0 | 1.881% | 0% | 0% | 1.881% | 43
Unintended System Motion | Endoscope/Camera | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Unintended System Motion | Endoscope/Camera | system | 2024 | Neither | 9 | 0 | 0 | 0.3354% | 0% | 0% | 0.3354% | 9
Unintended System Motion | Endoscope/Camera | system | 2025 | Neither | 15 | 0 | 0 | 0.4757% | 0% | 0% | 0.4757% | 15
Unintended System Motion | Endoscope/Camera | system | 2026 | Neither | 6 | 0 | 0 | 0.4567% | 0% | 0% | 0.4567% | 6
Unintended System Motion | Force Bipolar | energy | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Unintended System Motion | Force Bipolar | energy | 2023 | Neither | 13 | 0 | 0 | 0.5687% | 0% | 0% | 0.5687% | 13
Unintended System Motion | Force Bipolar | energy | 2024 | Neither | 18 | 0 | 0 | 0.6709% | 0% | 0% | 0.6709% | 18
Unintended System Motion | Force Bipolar | energy | 2025 | Neither | 24 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 24
Unintended System Motion | Force Bipolar | energy | 2026 | Neither | 10 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 10
Unintended System Motion | Cadiere Forceps | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Cadiere Forceps | mechanical | 2022 | Neither | 7 | 0 | 0 | 0.3733% | 0% | 0% | 0.3733% | 7
Unintended System Motion | Cadiere Forceps | mechanical | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Unintended System Motion | Cadiere Forceps | mechanical | 2024 | Neither | 12 | 0 | 0 | 0.4473% | 0% | 0% | 0.4473% | 12
Unintended System Motion | Cadiere Forceps | mechanical | 2025 | Neither | 20 | 0 | 0 | 0.6343% | 0% | 0% | 0.6343% | 20
Unintended System Motion | Cadiere Forceps | mechanical | 2026 | Neither | 13 | 0 | 0 | 0.9895% | 0% | 0% | 0.9895% | 13
Unintended System Motion | Permanent Cautery Hook | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Unintended System Motion | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Unintended System Motion | Permanent Cautery Hook | energy | 2023 | Neither | 15 | 0 | 0 | 0.6562% | 0% | 0% | 0.6562% | 15
Unintended System Motion | Permanent Cautery Hook | energy | 2024 | Neither | 16 | 0 | 0 | 0.5963% | 0% | 0% | 0.5963% | 16
Unintended System Motion | Permanent Cautery Hook | energy | 2025 | Neither | 19 | 0 | 0 | 0.6026% | 0% | 0% | 0.6026% | 19
Unintended System Motion | Permanent Cautery Hook | energy | 2026 | Neither | 4 | 0 | 0 | 0.3045% | 0% | 0% | 0.3045% | 4
Unintended System Motion | Tenaculum Forceps | mechanical | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Unintended System Motion | Tenaculum Forceps | mechanical | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Unintended System Motion | Tenaculum Forceps | mechanical | 2024 | Neither | 5 | 0 | 0 | 0.1864% | 0% | 0% | 0.1864% | 5
Unintended System Motion | Tenaculum Forceps | mechanical | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Unintended System Motion | Tenaculum Forceps | mechanical | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Unintended System Motion | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 6 | 0 | 0 | 0.2625% | 0% | 0% | 0.2625% | 6
Unintended System Motion | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Unintended System Motion | Tip-Up Fenestrated Grasper | mechanical | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Unintended System Motion | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Unintended System Motion | Bipolar (other) | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Unintended System Motion | Bipolar (other) | energy | 2023 | Neither | 4 | 0 | 0 | 0.175% | 0% | 0% | 0.175% | 4
Unintended System Motion | Bipolar (other) | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Unintended System Motion | Bipolar (other) | energy | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Unintended System Motion | Bipolar (other) | energy | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Unintended System Motion | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Unintended System Motion | SynchroSeal | energy | 2024 | Death | 0 | 1 | 0 | 0% | 0.03727% | 0% | 0.03727% | 1
Unintended System Motion | SynchroSeal | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Unintended System Motion | SynchroSeal | energy | 2025 | Neither | 9 | 0 | 0 | 0.2854% | 0% | 0% | 0.2854% | 9
Unintended System Motion | SynchroSeal | energy | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Unintended System Motion | Vision System | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Vision System | system | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Unintended System Motion | Vision System | system | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Unintended System Motion | Vision System | system | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Unintended System Motion | Vision System | system | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Unintended System Motion | Vision System | system | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Unintended System Motion | Permanent Cautery Spatula | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Unintended System Motion | Permanent Cautery Spatula | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Unintended System Motion | Permanent Cautery Spatula | energy | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Unintended System Motion | Permanent Cautery Spatula | energy | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Unintended System Motion | Permanent Cautery Spatula | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Unintended System Motion | Permanent Cautery Spatula | energy | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Unintended System Motion | Monopolar Cautery (other) | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Unintended System Motion | Monopolar Cautery (other) | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Unintended System Motion | Monopolar Cautery (other) | energy | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Unintended System Motion | Monopolar Cautery (other) | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Unintended System Motion | Harmonic (ultrasonic) | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Harmonic (ultrasonic) | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Unintended System Motion | Harmonic (ultrasonic) | energy | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Unintended System Motion | Generic System/Platform | system | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Unintended System Motion | Generic System/Platform | system | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Unintended System Motion | Generic System/Platform | system | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Unintended System Motion | Generic System/Platform | system | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Unintended System Motion | Generic System/Platform | system | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Unintended System Motion | Suction Irrigator | accessory | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Unintended System Motion | Suction Irrigator | accessory | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Unintended System Motion | Suction Irrigator | accessory | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Unintended System Motion | Suction Irrigator | accessory | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Unintended System Motion | Grasping Retractor | mechanical | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Unintended System Motion | Grasping Retractor | mechanical | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Unintended System Motion | Grasper/Forceps (other) | mechanical | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Unintended System Motion | Grasper/Forceps (other) | mechanical | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Unintended System Motion | Grasper/Forceps (other) | mechanical | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Unintended System Motion | Curved Bipolar Dissector | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Unintended System Motion | Curved Bipolar Dissector | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Unintended System Motion | PK Dissecting Forceps | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2021 | Neither | 141 | 0 | 0 | 8.846% | 0% | 0% | 8.846% | 141
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2022 | Neither | 139 | 0 | 0 | 7.413% | 0% | 0% | 7.413% | 139
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2023 | Injury | 0 | 0 | 2 | 0% | 0% | 0.08749% | 0.08749% | 2
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2023 | Neither | 176 | 0 | 0 | 7.699% | 0% | 0% | 7.699% | 176
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2024 | Neither | 261 | 0 | 0 | 9.728% | 0% | 0% | 9.728% | 261
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2025 | Injury | 0 | 0 | 9 | 0% | 0% | 0.2854% | 0.2854% | 9
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2025 | Neither | 7516 | 0 | 0 | 238.4% | 0% | 0% | 238.4% | 7516
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Material Split, Cut or Torn | Fenestrated Bipolar | energy | 2026 | Neither | 3182 | 0 | 0 | 242.2% | 0% | 0% | 242.2% | 3182
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2021 | Neither | 6 | 0 | 0 | 0.3764% | 0% | 0% | 0.3764% | 6
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2022 | Neither | 11 | 0 | 0 | 0.5867% | 0% | 0% | 0.5867% | 11
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2023 | Neither | 14 | 0 | 0 | 0.6124% | 0% | 0% | 0.6124% | 14
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2024 | Neither | 48 | 0 | 0 | 1.789% | 0% | 0% | 1.789% | 48
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2025 | Neither | 5670 | 0 | 0 | 179.8% | 0% | 0% | 179.8% | 5670
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Material Split, Cut or Torn | Monopolar Curved Scissors | energy | 2026 | Neither | 2918 | 0 | 0 | 222.1% | 0% | 0% | 222.1% | 2918
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2025 | Injury | 0 | 0 | 10 | 0% | 0% | 0.3172% | 0.3172% | 10
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2025 | Neither | 3747 | 0 | 0 | 118.8% | 0% | 0% | 118.8% | 3747
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2026 | Injury | 0 | 0 | 8 | 0% | 0% | 0.6089% | 0.6089% | 8
Material Split, Cut or Torn | ProGrasp Forceps | mechanical | 2026 | Neither | 1805 | 0 | 0 | 137.4% | 0% | 0% | 137.4% | 1805
Material Split, Cut or Torn | Maryland Bipolar | energy | 2021 | Neither | 54 | 0 | 0 | 3.388% | 0% | 0% | 3.388% | 54
Material Split, Cut or Torn | Maryland Bipolar | energy | 2022 | Neither | 65 | 0 | 0 | 3.467% | 0% | 0% | 3.467% | 65
Material Split, Cut or Torn | Maryland Bipolar | energy | 2023 | Neither | 97 | 0 | 0 | 4.243% | 0% | 0% | 4.243% | 97
Material Split, Cut or Torn | Maryland Bipolar | energy | 2024 | Neither | 108 | 0 | 0 | 4.025% | 0% | 0% | 4.025% | 108
Material Split, Cut or Torn | Maryland Bipolar | energy | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Material Split, Cut or Torn | Maryland Bipolar | energy | 2025 | Neither | 2998 | 0 | 0 | 95.08% | 0% | 0% | 95.08% | 2998
Material Split, Cut or Torn | Maryland Bipolar | energy | 2026 | Neither | 1056 | 0 | 0 | 80.38% | 0% | 0% | 80.38% | 1056
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2023 | Neither | 17 | 0 | 0 | 0.7437% | 0% | 0% | 0.7437% | 17
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2024 | Neither | 47 | 0 | 0 | 1.752% | 0% | 0% | 1.752% | 47
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2025 | Neither | 2368 | 0 | 0 | 75.1% | 0% | 0% | 75.1% | 2368
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2026 | Injury | 0 | 0 | 4 | 0% | 0% | 0.3045% | 0.3045% | 4
Material Split, Cut or Torn | Needle Driver / SutureCut | mechanical | 2026 | Neither | 1228 | 0 | 0 | 93.47% | 0% | 0% | 93.47% | 1228
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2025 | Injury | 0 | 0 | 4 | 0% | 0% | 0.1269% | 0.1269% | 4
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2025 | Neither | 1030 | 0 | 0 | 32.67% | 0% | 0% | 32.67% | 1030
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2026 | Injury | 0 | 0 | 3 | 0% | 0% | 0.2284% | 0.2284% | 3
Material Split, Cut or Torn | Cadiere Forceps | mechanical | 2026 | Neither | 499 | 0 | 0 | 37.98% | 0% | 0% | 37.98% | 499
Material Split, Cut or Torn | Force Bipolar | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Split, Cut or Torn | Force Bipolar | energy | 2021 | Neither | 27 | 0 | 0 | 1.694% | 0% | 0% | 1.694% | 27
Material Split, Cut or Torn | Force Bipolar | energy | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Split, Cut or Torn | Force Bipolar | energy | 2022 | Neither | 18 | 0 | 0 | 0.96% | 0% | 0% | 0.96% | 18
Material Split, Cut or Torn | Force Bipolar | energy | 2023 | Injury | 0 | 0 | 3 | 0% | 0% | 0.1312% | 0.1312% | 3
Material Split, Cut or Torn | Force Bipolar | energy | 2023 | Neither | 46 | 0 | 0 | 2.012% | 0% | 0% | 2.012% | 46
Material Split, Cut or Torn | Force Bipolar | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Split, Cut or Torn | Force Bipolar | energy | 2024 | Neither | 33 | 0 | 0 | 1.23% | 0% | 0% | 1.23% | 33
Material Split, Cut or Torn | Force Bipolar | energy | 2025 | Injury | 0 | 0 | 9 | 0% | 0% | 0.2854% | 0.2854% | 9
Material Split, Cut or Torn | Force Bipolar | energy | 2025 | Neither | 794 | 0 | 0 | 25.18% | 0% | 0% | 25.18% | 794
Material Split, Cut or Torn | Force Bipolar | energy | 2026 | Injury | 0 | 0 | 5 | 0% | 0% | 0.3806% | 0.3806% | 5
Material Split, Cut or Torn | Force Bipolar | energy | 2026 | Neither | 407 | 0 | 0 | 30.98% | 0% | 0% | 30.98% | 407
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2025 | Neither | 579 | 0 | 0 | 18.36% | 0% | 0% | 18.36% | 579
Material Split, Cut or Torn | Tip-Up Fenestrated Grasper | mechanical | 2026 | Neither | 359 | 0 | 0 | 27.33% | 0% | 0% | 27.33% | 359
Material Split, Cut or Torn | Bipolar (other) | energy | 2021 | Neither | 8 | 0 | 0 | 0.5019% | 0% | 0% | 0.5019% | 8
Material Split, Cut or Torn | Bipolar (other) | energy | 2022 | Neither | 9 | 0 | 0 | 0.48% | 0% | 0% | 0.48% | 9
Material Split, Cut or Torn | Bipolar (other) | energy | 2023 | Neither | 13 | 0 | 0 | 0.5687% | 0% | 0% | 0.5687% | 13
Material Split, Cut or Torn | Bipolar (other) | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Split, Cut or Torn | Bipolar (other) | energy | 2024 | Neither | 17 | 0 | 0 | 0.6336% | 0% | 0% | 0.6336% | 17
Material Split, Cut or Torn | Bipolar (other) | energy | 2025 | Injury | 0 | 0 | 5 | 0% | 0% | 0.1586% | 0.1586% | 5
Material Split, Cut or Torn | Bipolar (other) | energy | 2025 | Neither | 444 | 0 | 0 | 14.08% | 0% | 0% | 14.08% | 444
Material Split, Cut or Torn | Bipolar (other) | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Split, Cut or Torn | Bipolar (other) | energy | 2026 | Neither | 224 | 0 | 0 | 17.05% | 0% | 0% | 17.05% | 224
Material Split, Cut or Torn | Clip Applier | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Split, Cut or Torn | Clip Applier | mechanical | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Split, Cut or Torn | Clip Applier | mechanical | 2024 | Neither | 5 | 0 | 0 | 0.1864% | 0% | 0% | 0.1864% | 5
Material Split, Cut or Torn | Clip Applier | mechanical | 2025 | Neither | 455 | 0 | 0 | 14.43% | 0% | 0% | 14.43% | 455
Material Split, Cut or Torn | Clip Applier | mechanical | 2026 | Neither | 214 | 0 | 0 | 16.29% | 0% | 0% | 16.29% | 214
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2023 | Neither | 27 | 0 | 0 | 1.181% | 0% | 0% | 1.181% | 27
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2024 | Neither | 7 | 0 | 0 | 0.2609% | 0% | 0% | 0.2609% | 7
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2025 | Neither | 287 | 0 | 0 | 9.102% | 0% | 0% | 9.102% | 287
Material Split, Cut or Torn | Tenaculum Forceps | mechanical | 2026 | Neither | 76 | 0 | 0 | 5.785% | 0% | 0% | 5.785% | 76
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2021 | Injury | 0 | 0 | 1 | 0% | 0% | 0.06274% | 0.06274% | 1
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2021 | Neither | 15 | 0 | 0 | 0.941% | 0% | 0% | 0.941% | 15
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2022 | Neither | 13 | 0 | 0 | 0.6933% | 0% | 0% | 0.6933% | 13
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2023 | Neither | 13 | 0 | 0 | 0.5687% | 0% | 0% | 0.5687% | 13
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2024 | Neither | 21 | 0 | 0 | 0.7827% | 0% | 0% | 0.7827% | 21
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2025 | Injury | 0 | 0 | 3 | 0% | 0% | 0.09515% | 0.09515% | 3
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2025 | Neither | 234 | 0 | 0 | 7.422% | 0% | 0% | 7.422% | 234
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2026 | Injury | 0 | 0 | 1 | 0% | 0% | 0.07612% | 0.07612% | 1
Material Split, Cut or Torn | Permanent Cautery Hook | energy | 2026 | Neither | 91 | 0 | 0 | 6.927% | 0% | 0% | 6.927% | 91
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2021 | Neither | 4 | 0 | 0 | 0.2509% | 0% | 0% | 0.2509% | 4
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2022 | Neither | 3 | 0 | 0 | 0.16% | 0% | 0% | 0.16% | 3
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2023 | Neither | 25 | 0 | 0 | 1.094% | 0% | 0% | 1.094% | 25
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2024 | Neither | 9 | 0 | 0 | 0.3354% | 0% | 0% | 0.3354% | 9
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2025 | Neither | 208 | 0 | 0 | 6.597% | 0% | 0% | 6.597% | 208
Material Split, Cut or Torn | Grasping Retractor | mechanical | 2026 | Neither | 43 | 0 | 0 | 3.273% | 0% | 0% | 3.273% | 43
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2021 | Neither | 6 | 0 | 0 | 0.3764% | 0% | 0% | 0.3764% | 6
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2022 | Neither | 6 | 0 | 0 | 0.32% | 0% | 0% | 0.32% | 6
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2024 | Neither | 7 | 0 | 0 | 0.2609% | 0% | 0% | 0.2609% | 7
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2025 | Neither | 181 | 0 | 0 | 5.741% | 0% | 0% | 5.741% | 181
Material Split, Cut or Torn | Permanent Cautery Spatula | energy | 2026 | Neither | 82 | 0 | 0 | 6.242% | 0% | 0% | 6.242% | 82
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2021 | Injury | 0 | 0 | 6 | 0% | 0% | 0.3764% | 0.3764% | 6
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2021 | Neither | 16 | 0 | 0 | 1.004% | 0% | 0% | 1.004% | 16
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2022 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1067% | 0.1067% | 2
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2022 | Neither | 26 | 0 | 0 | 1.387% | 0% | 0% | 1.387% | 26
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2023 | Injury | 0 | 0 | 4 | 0% | 0% | 0.175% | 0.175% | 4
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2023 | Neither | 32 | 0 | 0 | 1.4% | 0% | 0% | 1.4% | 32
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2024 | Injury | 0 | 0 | 2 | 0% | 0% | 0.07454% | 0.07454% | 2
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2024 | Neither | 67 | 0 | 0 | 2.497% | 0% | 0% | 2.497% | 67
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2025 | Neither | 64 | 0 | 0 | 2.03% | 0% | 0% | 2.03% | 64
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Split, Cut or Torn | Tip Cover Accessory | energy | 2026 | Neither | 47 | 0 | 0 | 3.578% | 0% | 0% | 3.578% | 47
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2021 | Neither | 4 | 0 | 0 | 0.2509% | 0% | 0% | 0.2509% | 4
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2022 | Neither | 4 | 0 | 0 | 0.2133% | 0% | 0% | 0.2133% | 4
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2024 | Neither | 4 | 0 | 0 | 0.1491% | 0% | 0% | 0.1491% | 4
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2025 | Neither | 110 | 0 | 0 | 3.489% | 0% | 0% | 3.489% | 110
Material Split, Cut or Torn | Curved Bipolar Dissector | energy | 2026 | Neither | 45 | 0 | 0 | 3.425% | 0% | 0% | 3.425% | 45
Material Split, Cut or Torn | SynchroSeal | energy | 2023 | Injury | 0 | 0 | 1 | 0% | 0% | 0.04374% | 0.04374% | 1
Material Split, Cut or Torn | SynchroSeal | energy | 2023 | Neither | 8 | 0 | 0 | 0.35% | 0% | 0% | 0.35% | 8
Material Split, Cut or Torn | SynchroSeal | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Split, Cut or Torn | SynchroSeal | energy | 2024 | Neither | 29 | 0 | 0 | 1.081% | 0% | 0% | 1.081% | 29
Material Split, Cut or Torn | SynchroSeal | energy | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Split, Cut or Torn | SynchroSeal | energy | 2025 | Neither | 22 | 0 | 0 | 0.6977% | 0% | 0% | 0.6977% | 22
Material Split, Cut or Torn | SynchroSeal | energy | 2026 | Neither | 9 | 0 | 0 | 0.6851% | 0% | 0% | 0.6851% | 9
Material Split, Cut or Torn | Vessel Sealer | energy | 2021 | Neither | 6 | 0 | 0 | 0.3764% | 0% | 0% | 0.3764% | 6
Material Split, Cut or Torn | Vessel Sealer | energy | 2022 | Neither | 2 | 0 | 0 | 0.1067% | 0% | 0% | 0.1067% | 2
Material Split, Cut or Torn | Vessel Sealer | energy | 2023 | Neither | 7 | 0 | 0 | 0.3062% | 0% | 0% | 0.3062% | 7
Material Split, Cut or Torn | Vessel Sealer | energy | 2024 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03727% | 0.03727% | 1
Material Split, Cut or Torn | Vessel Sealer | energy | 2024 | Neither | 6 | 0 | 0 | 0.2236% | 0% | 0% | 0.2236% | 6
Material Split, Cut or Torn | Vessel Sealer | energy | 2025 | Neither | 17 | 0 | 0 | 0.5392% | 0% | 0% | 0.5392% | 17
Material Split, Cut or Torn | Vessel Sealer | energy | 2026 | Injury | 0 | 0 | 2 | 0% | 0% | 0.1522% | 0.1522% | 2
Material Split, Cut or Torn | Vessel Sealer | energy | 2026 | Neither | 11 | 0 | 0 | 0.8373% | 0% | 0% | 0.8373% | 11
Material Split, Cut or Torn | Grasper/Forceps (other) | mechanical | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Grasper/Forceps (other) | mechanical | 2025 | Neither | 26 | 0 | 0 | 0.8246% | 0% | 0% | 0.8246% | 26
Material Split, Cut or Torn | Grasper/Forceps (other) | mechanical | 2026 | Neither | 10 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 10
Material Split, Cut or Torn | Generic System/Platform | system | 2021 | Neither | 2 | 0 | 0 | 0.1255% | 0% | 0% | 0.1255% | 2
Material Split, Cut or Torn | Generic System/Platform | system | 2025 | Neither | 22 | 0 | 0 | 0.6977% | 0% | 0% | 0.6977% | 22
Material Split, Cut or Torn | Generic System/Platform | system | 2026 | Neither | 6 | 0 | 0 | 0.4567% | 0% | 0% | 0.4567% | 6
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Injury | 0 | 0 | 1 | 0% | 0% | 0.05333% | 0.05333% | 1
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2023 | Neither | 5 | 0 | 0 | 0.2187% | 0% | 0% | 0.2187% | 5
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2024 | Neither | 2 | 0 | 0 | 0.07454% | 0% | 0% | 0.07454% | 2
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Injury | 0 | 0 | 1 | 0% | 0% | 0.03172% | 0.03172% | 1
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Material Split, Cut or Torn | Stapler (SureForm/EndoWrist) | mechanical | 2026 | Neither | 10 | 0 | 0 | 0.7612% | 0% | 0% | 0.7612% | 10
Material Split, Cut or Torn | Patient-Side Cart/Arm | system | 2025 | Injury | 0 | 0 | 2 | 0% | 0% | 0.06343% | 0.06343% | 2
Material Split, Cut or Torn | Patient-Side Cart/Arm | system | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Material Split, Cut or Torn | Patient-Side Cart/Arm | system | 2026 | Neither | 19 | 0 | 0 | 1.446% | 0% | 0% | 1.446% | 19
Material Split, Cut or Torn | Cannula/Trocar/Reducer/Sheath | accessory | 2024 | Neither | 3 | 0 | 0 | 0.1118% | 0% | 0% | 0.1118% | 3
Material Split, Cut or Torn | Cannula/Trocar/Reducer/Sheath | accessory | 2025 | Neither | 6 | 0 | 0 | 0.1903% | 0% | 0% | 0.1903% | 6
Material Split, Cut or Torn | Vision System | system | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Split, Cut or Torn | Vision System | system | 2026 | Neither | 7 | 0 | 0 | 0.5328% | 0% | 0% | 0.5328% | 7
Material Split, Cut or Torn | Monopolar Cautery (other) | energy | 2023 | Neither | 2 | 0 | 0 | 0.08749% | 0% | 0% | 0.08749% | 2
Material Split, Cut or Torn | Monopolar Cautery (other) | energy | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Material Split, Cut or Torn | Monopolar Cautery (other) | energy | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Material Split, Cut or Torn | Endoscope/Camera | system | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Endoscope/Camera | system | 2025 | Neither | 4 | 0 | 0 | 0.1269% | 0% | 0% | 0.1269% | 4
Material Split, Cut or Torn | Endoscope/Camera | system | 2026 | Neither | 2 | 0 | 0 | 0.1522% | 0% | 0% | 0.1522% | 2
Material Split, Cut or Torn | Cobra Grasper | mechanical | 2025 | Neither | 2 | 0 | 0 | 0.06343% | 0% | 0% | 0.06343% | 2
Material Split, Cut or Torn | Cobra Grasper | mechanical | 2026 | Neither | 5 | 0 | 0 | 0.3806% | 0% | 0% | 0.3806% | 5
Material Split, Cut or Torn | Harmonic (ultrasonic) | energy | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Harmonic (ultrasonic) | energy | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Split, Cut or Torn | Harmonic (ultrasonic) | energy | 2025 | Neither | 3 | 0 | 0 | 0.09515% | 0% | 0% | 0.09515% | 3
Material Split, Cut or Torn | Harmonic (ultrasonic) | energy | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1
Material Split, Cut or Torn | PK Dissecting Forceps | energy | 2021 | Neither | 1 | 0 | 0 | 0.06274% | 0% | 0% | 0.06274% | 1
Material Split, Cut or Torn | PK Dissecting Forceps | energy | 2022 | Neither | 1 | 0 | 0 | 0.05333% | 0% | 0% | 0.05333% | 1
Material Split, Cut or Torn | PK Dissecting Forceps | energy | 2023 | Neither | 3 | 0 | 0 | 0.1312% | 0% | 0% | 0.1312% | 3
Material Split, Cut or Torn | Surgeon Console | system | 2023 | Neither | 1 | 0 | 0 | 0.04374% | 0% | 0% | 0.04374% | 1
Material Split, Cut or Torn | Surgeon Console | system | 2026 | Neither | 3 | 0 | 0 | 0.2284% | 0% | 0% | 0.2284% | 3
Material Split, Cut or Torn | Suction Irrigator | accessory | 2024 | Neither | 1 | 0 | 0 | 0.03727% | 0% | 0% | 0.03727% | 1
Material Split, Cut or Torn | Suction Irrigator | accessory | 2025 | Neither | 1 | 0 | 0 | 0.03172% | 0% | 0% | 0.03172% | 1
Material Split, Cut or Torn | Suction Irrigator | accessory | 2026 | Neither | 1 | 0 | 0 | 0.07612% | 0% | 0% | 0.07612% | 1